## Summary sheet
Lot 2707 | 1187/11.8nm BP
124mm Diameter X 5mm Thick
Delivered 124mm dia part was measured at center and then three spots near the edge of the filter to verify uniformity
Below information is taken from the measured delivered 124mm diameter part at 298K WITH the 65K correlation offset
added to ensure conformance with specifications at 65K
Lot 2707 | Center @65K | EdgeA @65K | EdgeB @65K | EdgeC @65K | EdgeD @65K | Specification | pass/fail?
Peak T% | 96.5897 | 96.692 | 96.4914 | 96.4191 | 96.6719 | > 90% | PASS
CWL: | 1186.5732 | 1186.0432 | 1186.3532 | 1186.3632 | 1186.4432 | 1187.0 +/- 2.3nm | PASS
FWHM: | 11.0895 | 11.0595 | 11.0595 | 11.0595 | 11.0495 | 11.8nm +/- 2.3nm | PASS

## Cold Test Data
Lot 2707
1187/11.8nm BP
1" cold test witness
This is for correlation purposes ONLY and is a witness from the same coating
run as the 124mm diameter part
| 298K | 77K | shift to 77K | calc shift 65K
peak T% | 97.5052 | 97.5398 | 0.034653 | 0.034653
CWL | 1184.69 | 1184.38 | -0.31 | -0.32683
FWHM | 10.83 | 10.82 | -0.01 | -0.010543
wl(nm) | T% | T%
1230 | -0.00086987 | 0.016741
1229.75 | 0.046406 | 0.041418
1229.5 | 0.028385 | 0.030145
1229.25 | 0.018501 | 0.050416
1229 | 0.016398 | 0.062944
1228.75 | 0.012723 | 0.015117
1228.5 | 0.013036 | 0.025572
1228.25 | 0.031247 | 0.073641
1228 | -0.0019929 | 0.010793
1227.75 | 0.0040793 | -0.0078648
1227.5 | 0.055033 | 0.093396
1227.25 | 0.036081 | 0.03863
1227 | 0.023612 | 0.021465
1226.75 | 0.053072 | 0.063528
1226.5 | 0.045472 | 0.064513
1226.25 | 0.026879 | 0.026331
1226 | 0.023468 | 0.059524
1225.75 | -0.0037691 | 0.031379
1225.5 | 0.036207 | 0.048839
1225.25 | 0.043561 | 0.060053
1225 | 0.041581 | 0.052643
1224.75 | 0.086876 | 0.046343
1224.5 | 0.038679 | 0.0026245
1224.25 | 0.04485 | 0.056337
1224 | 0.021854 | 0.033464
1223.75 | 0.035338 | 0.017613
1223.5 | 0.065188 | 0.052698
1223.25 | 0.021217 | 0.032173
1223 | 0.03561 | 0.051632
1222.75 | 0.02107 | 0.031748
1222.5 | 0.060264 | 0.021561
1222.25 | 0.022904 | 0.05522
1222 | 0.035611 | 0.020605
1221.75 | 0.044212 | 0.043886
1221.5 | 0.055549 | 0.087728
1221.25 | 0.010177 | 0.038423
1221 | 0.0066542 | 0.012465
1220.75 | 0.016235 | 0.024006
1220.5 | -0.00098427 | 0.022968
1220.25 | 0.011089 | 0.019742
1220 | 0.052831 | -0.012661
1219.75 | 0.04813 | 0.043394
1219.5 | -0.0014522 | 0.010164
1219.25 | 0.067339 | 0.064811
1219 | 0.039706 | 0.073532
1218.75 | 0.020918 | 0.031054
1218.5 | 0.035953 | 0.037044
1218.25 | 0.044099 | 0.028869
1218 | 0.019697 | 0.030554
1217.75 | 0.0039544 | 0.058585
1217.5 | 0.0007521 | 0.067814
1217.25 | 0.015338 | 0.031916
1217 | 0.017582 | 0.014748
1216.75 | 0.063338 | 0.030599
1216.5 | 0.034415 | 0.048496
1216.25 | 0.023458 | 0.040159
1216 | 0.030631 | 0.063803
1215.75 | -0.006999 | 0.024764
1215.5 | 0.083358 | 0.062616
1215.25 | 0.016366 | 0.062375
1215 | 0.027249 | 0.031278
1214.75 | 0.0062527 | 0.025231
1214.5 | 0.033482 | 0.044407
1214.25 | 0.017094 | 0.019434
1214 | 0.0066147 | 0.017977
1213.75 | -0.015076 | 0.0018028
1213.5 | 0.023127 | 0.035547
1213.25 | 0.044535 | 0.036279
1213 | 0.04428 | 0.013911
1212.75 | 0.051117 | 0.025346
1212.5 | 0.01679 | 0.019224
1212.25 | 0.0055574 | 0.054249
1212 | 0.011499 | 0.064857
1211.75 | 0.021876 | 0.034117
1211.5 | 0.04648 | 0.07259
1211.25 | 0.0093935 | 0.011432
1211 | 0.045271 | 0.062489
1210.75 | 0.029151 | 0.037888
1210.5 | 0.016705 | 0.030646
1210.25 | 0.036568 | 0.038354
1210 | 0.025205 | 0.057308
1209.75 | -0.0085383 | 0.010415
1209.5 | 0.021671 | 0.038738
1209.25 | 0.024593 | 0.066982
1209 | 0.012816 | 0.031963
1208.75 | 0.02457 | 0.048958
1208.5 | 0.05591 | 0.066317
1208.25 | 0.02461 | 0.027282
1208 | 0.027986 | 0.067837
1207.75 | 0.026801 | 0.043219
1207.5 | 0.0062638 | 0.050322
1207.25 | 0.038434 | 0.041479
1207 | -0.008369 | -0.018197
1206.75 | 0.038464 | 0.070353
1206.5 | -0.0092891 | 0.0010939
1206.25 | 0.029292 | 0.045445
1206 | 0.0029554 | 0.020231
1205.75 | 0.033687 | 0.085182
1205.5 | 0.014539 | 0.041948
1205.25 | 0.0034678 | 0.050618
1205 | 0.033953 | 0.064682
1204.75 | 0.029653 | 0.039936
1204.5 | 0.040193 | 0.043757
1204.25 | 0.046351 | 0.046242
1204 | 0.019419 | 0.053868
1203.75 | 0.045367 | 0.010275
1203.5 | 0.035372 | 0.059404
1203.25 | 0.039764 | 0.044893
1203 | 0.0010632 | 0.026973
1202.75 | 0.025766 | 0.069375
1202.5 | 0.032426 | 0.07587
1202.25 | 0.0098041 | 0.059909
1202 | -0.0035751 | 0.030681
1201.75 | 0.026307 | 0.033876
1201.5 | 0.021973 | 0.037098
1201.25 | 0.027241 | 0.053244
1201 | 0.056085 | 0.066515
1200.75 | 0.027691 | 0.054219
1200.5 | 0.026195 | 0.050037
1200.25 | 0.045601 | 0.064864
1200 | 0.034164 | 0.044289
1199.75 | 0.043823 | 0.022096
1199.5 | 0.064384 | 0.061311
1199.25 | 0.032463 | 0.061582
1199 | 0.066405 | 0.068246
1198.75 | 0.032364 | 0.050759
1198.5 | 0.042287 | 0.043564
1198.25 | 0.082071 | 0.089987
1198 | 0.065696 | 0.10061
1197.75 | 0.073997 | 0.062174
1197.5 | 0.084512 | 0.09587
1197.25 | 0.13728 | 0.13836
1197 | 0.12902 | 0.14197
1196.75 | 0.1421 | 0.12649
1196.5 | 0.17859 | 0.14545
1196.25 | 0.15679 | 0.16272
1196 | 0.2468 | 0.23998
1195.75 | 0.28702 | 0.24602
1195.5 | 0.30436 | 0.24726
1195.25 | 0.37519 | 0.31079
1195 | 0.44369 | 0.36574
1194.75 | 0.55507 | 0.4126
1194.5 | 0.71036 | 0.56768
1194.25 | 0.84333 | 0.62194
1194 | 1.0556 | 0.84296
1193.75 | 1.2808 | 0.97194
1193.5 | 1.6329 | 1.2247
1193.25 | 2.0574 | 1.5321
1193 | 2.5587 | 1.9264
1192.75 | 3.2861 | 2.4227
1192.5 | 4.4075 | 3.2343
1192.25 | 5.6495 | 4.0872
1192 | 7.4867 | 5.3934
1191.75 | 9.7226 | 7.0106
1191.5 | 13.138 | 9.4121
1191.25 | 16.9921 | 12.2427
1191 | 21.5861 | 15.6478
1190.75 | 28.987 | 21.505
1190.5 | 35.7961 | 27.099
1190.25 | 43.835 | 34.0137
1190 | 52.1778 | 41.7447
1189.75 | 60.7345 | 50.3415
1189.5 | 68.3931 | 58.6101
1189.25 | 74.6075 | 65.7843
1189 | 81.1618 | 73.9221
1188.75 | 85.169 | 79.1981
1188.5 | 88.2158 | 83.5984
1188.25 | 90.5596 | 87.1591
1188 | 92.1089 | 89.6894
1187.75 | 93.2111 | 91.4962
1187.5 | 93.9425 | 92.6373
1187.25 | 94.6028 | 93.7451
1187 | 95.0799 | 94.3992
1186.75 | 95.4729 | 94.9416
1186.5 | 95.8613 | 95.4648
1186.25 | 96.2176 | 95.9562
1186 | 96.482 | 96.301
1185.75 | 96.7361 | 96.6014
1185.5 | 96.9811 | 96.8841
1185.25 | 97.0218 | 96.9862
1185 | 97.1252 | 97.1123
1184.75 | 97.2004 | 97.1908
1184.5 | 97.1971 | 97.1873
1184.25 | 97.2579 | 97.1975
1184 | 97.3436 | 97.2075
1183.75 | 97.3373 | 97.2301
1183.5 | 97.4109 | 97.2903
1183.25 | 97.4831 | 97.351
1183 | 97.5052 | 97.4462
1182.75 | 97.4828 | 97.4856
1182.5 | 97.3361 | 97.5398
1182.25 | 97.0408 | 97.4989
1182 | 96.4236 | 97.2234
1181.75 | 95.4551 | 96.761
1181.5 | 94.2395 | 96.1117
1181.25 | 91.8655 | 94.6216
1181 | 89.2606 | 92.8905
1180.75 | 85.6928 | 90.3974
1180.5 | 81.1536 | 87.0186
1180.25 | 75.5357 | 82.4115
1180 | 69.3778 | 77.268
1179.75 | 63.1958 | 71.8029
1179.5 | 54.5785 | 63.6527
1179.25 | 47.9036 | 57.0849
1179 | 40.6995 | 49.7109
1178.75 | 33.7876 | 42.14
1178.5 | 27.7345 | 35.2539
1178.25 | 22.412 | 28.9635
1178 | 18.3196 | 23.9381
1177.75 | 13.8415 | 18.3443
1177.5 | 11.1084 | 14.8569
1177.25 | 8.8249 | 11.825
1177 | 6.8028 | 9.1158
1176.75 | 5.293 | 7.1452
1176.5 | 4.1499 | 5.5658
1176.25 | 3.3015 | 4.3849
1176 | 2.519 | 3.3001
1175.75 | 2.0241 | 2.6249
1175.5 | 1.579 | 2.0774
1175.25 | 1.2841 | 1.657
1175 | 1.0251 | 1.3566
1174.75 | 0.82244 | 1.0713
1174.5 | 0.68564 | 0.86215
1174.25 | 0.62878 | 0.71623
1174 | 0.48071 | 0.57017
1173.75 | 0.41614 | 0.50403
1173.5 | 0.2855 | 0.39062
1173.25 | 0.29478 | 0.33349
1173 | 0.2071 | 0.25505
1172.75 | 0.21168 | 0.28131
1172.5 | 0.17037 | 0.21648
1172.25 | 0.15257 | 0.19622
1172 | 0.13143 | 0.16064
1171.75 | 0.10195 | 0.11007
1171.5 | 0.10929 | 0.1008
1171.25 | 0.068641 | 0.077747
1171 | 0.093697 | 0.093805
1170.75 | 0.064231 | 0.13658
1170.5 | 0.056351 | 0.077797
1170.25 | 0.048073 | 0.065499
1170 | 0.096737 | 0.083738
1169.75 | 0.03937 | 0.017676
1169.5 | 0.061845 | 0.056041
1169.25 | 0.013403 | 0.037196
1169 | 0.025797 | 0.066962
1168.75 | 0.071476 | 0.079626
1168.5 | 0.017715 | 0.026245
1168.25 | 0.056093 | 0.042081
1168 | 0.06701 | 0.03992
1167.75 | 0.059747 | 0.045066
1167.5 | 0.034601 | 0.046776
1167.25 | 0.0021758 | 0.020835
1167 | 0.025994 | 0.040767
1166.75 | 0.020162 | 0.037016
1166.5 | 0.028375 | 0.038441
1166.25 | -0.0089198 | 0.030505
1166 | 0.052714 | 0.047251
1165.75 | 0.012808 | 0.022131
1165.5 | 0.035754 | 0.017096
1165.25 | 0.047217 | 0.061164
1165 | 0.038674 | 0.026498
1164.75 | -0.0057035 | 0.0076661
1164.5 | 0.014705 | 0.015699
1164.25 | 0.063879 | 0.042924
1164 | 0.040755 | 0.031029
1163.75 | 1.4292e-05 | 0.023243
1163.5 | 0.027319 | 0.021099
1163.25 | 0.052692 | 0.046583
1163 | 0.036238 | 0.032199
1162.75 | 0.025705 | 0.031046
1162.5 | -0.0035591 | 0.026575
1162.25 | 0.08289 | 0.02048
1162 | -0.0057763 | 0.019186
1161.75 | 0.075718 | 0.044885
1161.5 | 0.048627 | 0.068761
1161.25 | 0.045096 | 0.064139
1161 | 0.0067931 | 0.023606
1160.75 | 0.014506 | 0.017093
1160.5 | 0.027103 | 0.021822
1160.25 | 0.084755 | 0.051717
1160 | 0.019787 | 0.027244
1159.75 | 0.058246 | 0.076712
1159.5 | 0.048011 | 0.04345
1159.25 | 0.03832 | 0.047206
1159 | 0.047433 | 0.039315
1158.75 | 0.036118 | 0.044944
1158.5 | 0.011127 | 0.064926
1158.25 | -0.0014793 | 0.017021
1158 | 0.018268 | 0.035594
1157.75 | 0.0039045 | 0.05386
1157.5 | 0.034902 | 0.060026
1157.25 | 0.014624 | 0.033294
1157 | 0.04391 | 0.081556
1156.75 | 0.0011883 | -0.007539
1156.5 | 0.040227 | 0.055146
1156.25 | 0.01707 | 0.037557
1156 | 0.0489 | 0.069213
1155.75 | 0.048811 | 0.062997
1155.5 | 0.096253 | 0.099112
1155.25 | 0.054416 | 0.011467
1155 | 0.0023753 | 0.06321
1154.75 | 0.028149 | 0.039914
1154.5 | 0.037432 | 0.025049
1154.25 | 0.075877 | 0.056698
1154 | -0.0082627 | 0.027348
1153.75 | 0.04165 | 0.07061
1153.5 | 0.024196 | 0.032743
1153.25 | 0.064196 | 0.049242
1153 | 0.043585 | 0.063453
1152.75 | 0.03234 | 0.030082
1152.5 | 0.035923 | 0.058178
1152.25 | 0.0051852 | 0.031964
1152 | 0.011885 | 0.017081
1151.75 | 0.022156 | 0.025335
1151.5 | 0.0050978 | 0.034498
1151.25 | -0.0085762 | 0.021644
1151 | 0.070644 | 0.032658
1150.75 | 0.076798 | 0.042446
1150.5 | 0.044252 | 0.042605
1150.25 | 0.034148 | 0.054261
1150 | 0.043404 | 0.079014

## T% and BLKG
Lot 2707 | 1187/11.8nm BP
Information in red is the Room temp (298K) measured data for the delivered 124mm diameter filter
Information in black is the calculated (assumed) shift of the 124mm diameter filter at 65K
Lot 2707
Lot 2707 | [10 empty] | 1187/11.8nm
124mm Dia X 5mm Thick | [10 empty] | Out-of-band Blocking
Measured at Center of part |  |  | Edge A |  | Edge B |  | Edge C |  | Edge D |  | Measured from 1000-1600nm ONLY
| 298K | @65K | 298K | @65K | 298K | @65K | 298K | @65K | 298K | @65K | Will be OD4 from 200-6000nm when used in series
peakT% | 96.555 | 96.5897 | 96.6574 | 96.692 | 96.4567 | 96.4914 | 96.3845 | 96.4191 | 96.6372 | 96.6719 | with the MKO J filter from same order
CWL | 1186.9 | 1186.5732 | 1186.37 | 1186.0432 | 1186.68 | 1186.3532 | 1186.69 | 1186.3632 | 1186.77 | 1186.4432
FWHM | 11.1 | 11.0895 | 11.07 | 11.0595 | 11.07 | 11.0595 | 11.07 | 11.0595 | 11.06 | 11.0495
wl(nm) | T% |  | T% |  | T% |  | T% |  | T% |  | wl(nm) | OD
1230 | 0.0068872 |  | 0.0012168 |  | -0.0048799 |  | -0.0073085 |  | -0.0036585 |  | 1600 | 5.4546
1229.9 | -0.01887 |  | -0.0075541 |  | -0.011344 |  | 0.002516 |  | -0.012596 |  | 1599.5 | 5.4605
1229.8 | 0 |  | -0.0050349 |  | -0.0037748 |  | 0 |  | -0.0050419 |  | 1599 | 5.4623
1229.7 | 0 |  | -0.0012214 |  | -0.010938 |  | 0 |  | -0.012211 |  | 1598.5 | 5.464
1229.6 | 0.0036544 |  | 0.0048583 |  | 0.0036514 |  | 0.0024414 |  | -0.007304 |  | 1598 | 5.4771
1229.5 | 0 |  | 0.005034 |  | -0.010076 |  | 0 |  | -0.012596 |  | 1597.5 | 5.4824
1229.4 | -0.0037731 |  | 0.0025178 |  | -0.0063021 |  | -0.0012592 |  | -0.0075621 |  | 1597 | 5.4858
1229.3 | 0.0048798 |  | 0.0024402 |  | 2.2213e-07 |  | 0.0024404 |  | -0.011008 |  | 1596.5 | 5.4921
1229.2 | -0.0024407 |  | -0.0036454 |  | -0.0012168 |  | 0.0036628 |  | -0.0072989 |  | 1596 | 5.4984
1229.1 | -0.0012553 |  | -0.0012568 |  | -0.0050398 |  | 0.0037704 |  | -0.0075567 |  | 1595.5 | 5.5038
1229 | -0.0025124 |  | 0.0012589 |  | -0.0088242 |  | -0.0037699 |  | 0 |  | 1595 | 5.5033
1228.9 | -0.00366 |  | -0.0036644 |  | -0.0024442 |  | -0.0012198 |  | -0.011005 |  | 1594.5 | 5.5162
1228.8 | -0.0073195 |  | 0.0036432 |  | 0.0048658 |  | 0.0024389 |  | -0.0024338 |  | 1594 | 5.5197
1228.7 | -0.0025082 |  | -0.001256 |  | -0.0062956 |  | -0.013833 |  | -0.0075541 |  | 1593.5 | 5.5219
1228.6 | -0.010047 |  | -0.0012578 |  | -0.0037804 |  | 0.0050341 |  | -0.015129 |  | 1593 | 5.5354
1228.5 | -0.0036564 |  | 0 |  | -0.003665 |  | 0 |  | -0.0061148 |  | 1592.5 | 5.5282
1228.4 | -0.0024428 |  | 0.0036704 |  | -0.0036775 |  | -0.0024432 |  | -0.0036514 |  | 1592 | 5.5459
1228.3 | -0.010024 |  | -0.0062679 |  | -0.0025133 |  | -0.015095 |  | -0.010075 |  | 1591.5 | 5.5488
1228.2 | -0.0050189 |  | -0.0012555 |  | -0.010065 |  | 0 |  | -0.016388 |  | 1591 | 5.5485
1228.1 | -0.007309 |  | 0 |  | -0.0048857 |  | -2.2141e-07 |  | -0.0073297 |  | 1590.5 | 5.5585
1228 | -0.002444 |  | -0.004891 |  | -0.0085689 |  | 0.0012202 |  | -0.0061291 |  | 1590 | 5.5656
1227.9 | -0.0075202 |  | -0.0087783 |  | -0.0012564 |  | -0.011327 |  | -0.012573 |  | 1589.5 | 5.564
1227.8 | -0.0075311 |  | -0.0062781 |  | 0.0037757 |  | -0.0025159 |  | -0.0088146 |  | 1589 | 5.5817
1227.7 | -0.0012175 |  | -0.007316 |  | 0.0024401 |  | -0.0048623 |  | -0.0073227 |  | 1588.5 | 5.5766
1227.6 | -0.01343 |  | 0.0012238 |  | -0.0073043 |  | -0.0012205 |  | -0.012232 |  | 1588 | 5.5855
1227.5 | -0.010074 |  | 0.0012558 |  | -0.0012584 |  | -0.0088122 |  | -0.010059 |  | 1587.5 | 5.5903
1227.4 | -0.0025158 |  | 0.0075414 |  | 0.0037769 |  | 0.0075407 |  | -0.0075545 |  | 1587 | 5.6
1227.3 | -0.012164 |  | -0.0012198 |  | -0.0085521 |  | -0.0012153 |  | -0.0085424 |  | 1586.5 | 5.5981
1227.2 | -0.0073182 |  | 0 |  | -0.0036736 |  | -0.003655 |  | -0.0048654 |  | 1586 | 5.6097
1227.1 | -0.010072 |  | 0.0012565 |  | 0.0062907 |  | -0.0088114 |  | -0.0037769 |  | 1585.5 | 5.6108
1227 | -0.011326 |  | -0.001258 |  | 0 |  | -0.010061 |  | -0.012609 |  | 1585 | 5.627
1226.9 | -0.0048606 |  | 0.0012187 |  | 0.0012202 |  | -0.015788 |  | -0.010985 |  | 1584.5 | 5.6253
1226.8 | -0.0036587 |  | -0.0024427 |  | -0.0024474 |  | -0.018296 |  | -0.0097402 |  | 1584 | 5.6266
1226.7 | -0.011334 |  | 0 |  | -0.0062835 |  | -0.015127 |  | -0.012594 |  | 1583.5 | 5.6287
1226.6 | -0.010058 |  | 0.007546 |  | -0.0088012 |  | -0.016349 |  | -0.016386 |  | 1583 | 5.642
1226.5 | -0.013363 |  | 0.0060945 |  | -0.010968 |  | -0.019525 |  | -0.015876 |  | 1582.5 | 5.6447
1226.4 | -0.0073087 |  | 0.0024395 |  | -0.0097594 |  | -0.017069 |  | -0.0073054 |  | 1582 | 5.6429
1226.3 | -0.011326 |  | -0.0063013 |  | -0.0063014 |  | -0.0037798 |  | -0.0062994 |  | 1581.5 | 5.6557
1226.2 | -0.0037727 |  | 0 |  | -0.003778 |  | -0.008809 |  | -0.0088312 |  | 1581 | 5.656
1226.1 | -0.017022 |  | -0.0048756 |  | -0.0097425 |  | -0.010995 |  | -0.010952 |  | 1580.5 | 5.6641
1226 | -0.015827 |  | -0.0012206 |  | -0.0024417 |  | 2.2205e-07 |  | -0.006097 |  | 1580 | 5.6761
1225.9 | -0.01384 |  | -0.0062981 |  | 0.0012585 |  | -0.0025172 |  | -0.011325 |  | 1579.5 | 5.6801
1225.8 | -0.018882 |  | 0.0050364 |  | -0.011318 |  | -0.0012592 |  | -0.012554 |  | 1579 | 5.6782
1225.7 | -0.0061152 |  | 0.0097512 |  | -0.0097491 |  | -0.0085504 |  | -0.013417 |  | 1578.5 | 5.685
1225.6 | -0.013387 |  | -0.0048865 |  | -0.0024404 |  | -0.0097754 |  | -0.015874 |  | 1578 | 5.6859
1225.5 | -0.016364 |  | -0.0063082 |  | -0.0075471 |  | -0.017582 |  | -0.013848 |  | 1577.5 | 5.6956
1225.4 | -0.0075565 |  | 0.0037778 |  | 0.0037756 |  | -0.012565 |  | -0.010036 |  | 1577 | 5.6878
1225.3 | -0.011003 |  | 0 |  | -0.0061132 |  | -0.010969 |  | -0.0048786 |  | 1576.5 | 5.6987
1225.2 | -0.0085207 |  | -0.0036626 |  | -0.0060918 |  | -0.010995 |  | -0.017102 |  | 1576 | 5.7012
1225.1 | -0.0088134 |  | -0.0025221 |  | -0.0075513 |  | -0.013812 |  | -0.008813 |  | 1575.5 | 5.7038
1225 | -0.021402 |  | 0.0050353 |  | -0.0012585 |  | -0.0037695 |  | -0.0087677 |  | 1575 | 5.7112
1224.9 | -0.010997 |  | -0.003659 |  | -0.013426 |  | -0.014637 |  | -0.0060976 |  | 1574.5 | 5.7067
1224.8 | -0.010956 |  | -0.0036614 |  | -0.0072984 |  | -0.0073353 |  | -0.0073184 |  | 1574 | 5.7205
1224.7 | -0.015095 |  | -0.0063028 |  | 0.0012567 |  | -0.0075312 |  | -0.01638 |  | 1573.5 | 5.7245
1224.6 | -0.0088148 |  | 0.0075692 |  | 0.0012578 |  | -0.0037721 |  | -0.010086 |  | 1573 | 5.7259
1224.5 | -0.0061033 |  | 0 |  | -0.010975 |  | -0.0060967 |  | -0.017114 |  | 1572.5 | 5.73
1224.4 | -0.0085197 |  | 0.0036605 |  | -0.0097861 |  | -0.0085493 |  | -0.0085386 |  | 1572 | 5.7351
1224.3 | -0.010063 |  | -0.003783 |  | 0.0050222 |  | 0 |  | -0.0062967 |  | 1571.5 | 5.7336
1224.2 | -0.010074 |  | 0 |  | 0.010055 |  | -0.0062811 |  | -0.0025109 |  | 1571 | 5.7452
1224.1 | -0.013429 |  | 0.0060978 |  | -0.0048817 |  | -0.0036547 |  | -0.015806 |  | 1570.5 | 5.7388
1224 | -0.0085267 |  | 0.0060981 |  | -0.003666 |  | -0.0085506 |  | -0.010975 |  | 1570 | 5.7541
1223.9 | -0.021382 |  | 0 |  | -0.0025105 |  | -0.010033 |  | -0.011245 |  | 1569.5 | 5.7504
1223.8 | -0.013855 |  | 0.0050373 |  | 0.0075456 |  | -0.011314 |  | -0.011269 |  | 1569 | 5.7559
1223.7 | -0.0012206 |  | 0.0061072 |  | -0.0024419 |  | -0.0012187 |  | -0.0097123 |  | 1568.5 | 5.7665
1223.6 | -0.0085265 |  | 0.0024411 |  | -0.008566 |  | -0.0061049 |  | -0.013407 |  | 1568 | 5.757
1223.5 | -0.016346 |  | -0.0037787 |  | -0.011302 |  | -0.012609 |  | -0.0025088 |  | 1567.5 | 5.7644
1223.4 | -0.01133 |  | 0.0037825 |  | -0.0050333 |  | -0.0088259 |  | -0.0087948 |  | 1567 | 5.7599
1223.3 | -0.00977 |  | 0.0048873 |  | -0.0012216 |  | -0.010977 |  | -0.01218 |  | 1566.5 | 5.7786
1223.2 | -0.017138 |  | 0.0061028 |  | -0.007339 |  | 0.0024432 |  | -0.0036661 |  | 1566 | 5.7768
1223.1 | -0.0075392 |  | 0.0025189 |  | 0.0050241 |  | -0.0037844 |  | -0.011076 |  | 1565.5 | 5.7783
1223 | -0.0062905 |  | 0.0037795 |  | -0.0012575 |  | 0.0012608 |  | -0.0050387 |  | 1565 | 5.7891
1222.9 | -0.015863 |  | 0 |  | -0.010983 |  | -0.013399 |  | -0.010948 |  | 1564.5 | 5.7792
1222.8 | -0.0085617 |  | 0.0024396 |  | -0.00856 |  | -0.0036637 |  | -0.014679 |  | 1564 | 5.7782
1222.7 | -0.010053 |  | 0.0075576 |  | -2.2855e-07 |  | 0.0025175 |  | -0.009918 |  | 1563.5 | 5.7922
1222.6 | -0.0050328 |  | 0.0025197 |  | 0.0088041 |  | -0.0012575 |  | 0.0012636 |  | 1563 | 5.7927
1222.5 | -0.018309 |  | 0.0036616 |  | -0.0036583 |  | -0.0012198 |  | -0.004868 |  | 1562.5 | 5.7928
1222.4 | -0.019553 |  | -0.0085303 |  | 0.0012213 |  | 0.010938 |  | -0.0024498 |  | 1562 | 5.8122
1222.3 | -0.0050322 |  | -0.0075572 |  | 0 |  | -0.0075686 |  | -0.003714 |  | 1561.5 | 5.7887
1222.2 | -0.012595 |  | -0.0037829 |  | 0.0063012 |  | -0.0037807 |  | -0.012373 |  | 1561 | 5.8033
1222.1 | -0.014636 |  | -0.0024409 |  | 0.0060873 |  | 0.0072942 |  | -0.013601 |  | 1560.5 | 5.8252
1222 | -0.01465 |  | -0.002438 |  | -0.003662 |  | -0.0048669 |  | -0.0049465 |  | 1560 | 5.8205
1221.9 | -0.0037847 |  | -0.0012586 |  | -0.0037811 |  | -0.0025188 |  | -0.0061862 |  | 1559.5 | 5.8278
1221.8 | 0 |  | 0.0063046 |  | -0.0037789 |  | 0 |  | -0.0037121 |  | 1559 | 5.8239
1221.7 | -0.018295 |  | 0.0024412 |  | 0.0012182 |  | 0.0024286 |  | -0.004948 |  | 1558.5 | 5.8154
1221.6 | -0.010981 |  | 0.0036664 |  | 0.0048808 |  | -0.002435 |  | -0.016077 |  | 1558 | 5.8236
1221.5 | -0.013848 |  | 0.0050288 |  | -0.0012603 |  | -0.0063032 |  | -0.0037131 |  | 1557.5 | 5.836
1221.4 | -0.0075553 |  | 0.0037807 |  | -0.0012592 |  | 0.0012557 |  | 0.0024769 |  | 1557 | 5.8327
1221.3 | -0.0060737 |  | 0.0036577 |  | 0.0012184 |  | 0.002433 |  | -0.012374 |  | 1556.5 | 5.8329
1221.2 | 0.0012156 |  | 0.0060839 |  | 0.0048777 |  | -0.0073132 |  | -0.0099018 |  | 1556 | 5.8391
1221.1 | -0.0025134 |  | 0.011313 |  | 0.0012587 |  | -0.006295 |  | -0.018475 |  | 1555.5 | 5.8399
1221 | -0.0088196 |  | 0.0012598 |  | 0.0062835 |  | -0.0037555 |  | -0.0048942 |  | 1555 | 5.8602
1220.9 | 0.0012118 |  | 0.0073106 |  | 0 |  | 0.0024245 |  | -0.0061687 |  | 1554.5 | 5.8464
1220.8 | 0.0024327 |  | 0.0036492 |  | 0.0085361 |  | 0 |  | -0.0050522 |  | 1554 | 5.8505
1220.7 | -0.013837 |  | 0.0062779 |  | -0.0012592 |  | 0.013814 |  | 0 |  | 1553.5 | 5.8494
1220.6 | -0.011341 |  | 0.0037761 |  | -0.0037822 |  | -0.0012587 |  | -0.0036639 |  | 1553 | 5.8465
1220.5 | 0.0012116 |  | 0.001221 |  | 0.0073338 |  | 0 |  | -0.012374 |  | 1552.5 | 5.8604
1220.4 | -0.011027 |  | 0.0048703 |  | 0.0061012 |  | -0.0048937 |  | -0.008667 |  | 1552 | 5.8634
1220.3 | -0.0099372 |  | 0.0037657 |  | 0.0037776 |  | 0.011328 |  | -0.0037141 |  | 1551.5 | 5.8638
1220.2 | -0.01115 |  | -0.010069 |  | -0.0012606 |  | 0.0049365 |  | -0.0037126 |  | 1551 | 5.8541
1220.1 | -0.012228 |  | 0.0036498 |  | 0.002433 |  | 0.0024664 |  | -0.00738 |  | 1550.5 | 5.8579
1220 | -0.0098048 |  | -0.0049041 |  | 0 |  | 0.0061032 |  | -0.007542 |  | 1550 | 5.8805
1219.9 | -0.0074858 |  | 0.0061624 |  | 0.0050392 |  | 0.0012617 |  | -0.004827 |  | 1549.5 | 5.873
1219.8 | -0.010996 |  | -0.0012648 |  | 0.0050433 |  | -0.00251 |  | -0.0050659 |  | 1549 | 5.8696
1219.7 | -0.0099894 |  | -0.0085384 |  | 0.0060777 |  | -0.0060734 |  | -0.014635 |  | 1548.5 | 5.8794
1219.6 | -0.0088282 |  | 0.0048867 |  | 0.0060949 |  | 0.0048819 |  | -0.013483 |  | 1548 | 5.9016
1219.5 | -0.002438 |  | -0.010041 |  | 0.0025191 |  | -0.0025159 |  | -0.0062706 |  | 1547.5 | 5.8852
1219.4 | -0.0097665 |  | 0.0062805 |  | 0.0037838 |  | -0.0025194 |  | -0.0036474 |  | 1547 | 5.8754
1219.3 | -0.011262 |  | -0.0097229 |  | -0.0024344 |  | 0.0012197 |  | -0.010037 |  | 1546.5 | 5.9032
1219.2 | -0.0024458 |  | 0 |  | 0.01092 |  | 0.013456 |  | -0.013831 |  | 1546 | 5.8905
1219.1 | -0.0012358 |  | -0.008805 |  | 0.0012579 |  | 0.0037005 |  | -0.012209 |  | 1545.5 | 5.9018
1219 | -0.012271 |  | 0 |  | 0.010049 |  | 0.0050601 |  | -0.0061151 |  | 1545 | 5.9019
1218.9 | -0.0074117 |  | -0.0012145 |  | 0.004882 |  | -0.0036439 |  | 0.0012562 |  | 1544.5 | 5.9064
1218.8 | 2.3015e-07 |  | -0.0072916 |  | 0.0012212 |  | -0.0012303 |  | -0.010962 |  | 1544 | 5.9116
1218.7 | -0.003661 |  | -0.0075719 |  | 0.0025162 |  | 0.0049642 |  | -0.018806 |  | 1543.5 | 5.9093
1218.6 | -0.0024414 |  | -0.0088133 |  | 0.0075328 |  | 0.0012208 |  | -0.0086503 |  | 1543 | 5.9067
1218.5 | -0.0088293 |  | -0.0036681 |  | 0.0036638 |  | 0.0012587 |  | -0.0074141 |  | 1542.5 | 5.9059
1218.4 | -0.0088366 |  | 0.0024378 |  | 0.01215 |  | 0.0037061 |  | -0.0097632 |  | 1542 | 5.9225
1218.3 | -0.0085598 |  | -0.01383 |  | -0.0050117 |  | -0.0012359 |  | -0.012564 |  | 1541.5 | 5.9192
1218.2 | 0 |  | -0.0075681 |  | 0.0012557 |  | 0.0097752 |  | -0.010951 |  | 1541 | 5.9186
1218.1 | -0.0050176 |  | -0.0060707 |  | 0.0060802 |  | 0.0062537 |  | -0.0025092 |  | 1540.5 | 5.9239
1218 | 0.0073693 |  | 0 |  | 0.0048829 |  | 0.013492 |  | 0 |  | 1540 | 5.9171
1217.9 | 0.0024657 |  | -0.0088222 |  | 0 |  | 0.003719 |  | -0.013442 |  | 1539.5 | 5.9315
1217.8 | 0.0085306 |  | -0.010042 |  | 0.0036929 |  | 0.0036905 |  | -0.0024448 |  | 1539 | 5.9245
1217.7 | 0.0025181 |  | -0.0085339 |  | -0.0012332 |  | -0.0049684 |  | 0 |  | 1538.5 | 5.9343
1217.6 | -0.0037502 |  | -0.0085792 |  | 0.010968 |  | -0.0024495 |  | -0.0097993 |  | 1538 | 5.9376
1217.5 | -0.0048399 |  | -0.0037022 |  | -0.0012614 |  | 0.0037818 |  | -0.0037127 |  | 1537.5 | 5.9384
1217.4 | -0.0012185 |  | -0.003789 |  | 0.0025068 |  | -0.0062634 |  | -0.0024576 |  | 1537 | 5.9356
1217.3 | -0.0050303 |  | -0.013283 |  | 0.012229 |  | -0.0060671 |  | -0.0061913 |  | 1536.5 | 5.9473
1217.2 | 0.0049428 |  | -0.010179 |  | 0.0036817 |  | 0.0012259 |  | -0.011004 |  | 1536 | 5.9459
1217.1 | 0.0012353 |  | -0.014608 |  | 0.0086507 |  | 0.0025205 |  | -0.007553 |  | 1535.5 | 5.9426
1217 | 0.0048887 |  | -0.0085439 |  | 0.0050638 |  | 0 |  | -0.010016 |  | 1535 | 5.9511
1216.9 | 0 |  | -0.0088192 |  | -0.0073001 |  | -0.013891 |  | -0.0036353 |  | 1534.5 | 5.9481
1216.8 | -0.0036745 |  | -0.011271 |  | -0.0061069 |  | -0.0024986 |  | -0.0073136 |  | 1534 | 5.9546
1216.7 | 0.0037137 |  | -0.0072984 |  | -0.0049193 |  | 0.0012122 |  | -0.0050364 |  | 1533.5 | 5.9506
1216.6 | 0.006151 |  | -0.0036976 |  | 0 |  | -0.0049078 |  | -0.014887 |  | 1533 | 5.9486
1216.5 | 0.0024818 |  | -0.011195 |  | -0.003645 |  | -0.0075542 |  | -0.0061154 |  | 1532.5 | 5.9685
1216.4 | 0.0061162 |  | -0.0085757 |  | -0.0012198 |  | -0.010914 |  | 0 |  | 1532 | 5.9566
1216.3 | -0.0062984 |  | -0.0050408 |  | 0.0025168 |  | -0.0012627 |  | -0.0049837 |  | 1531.5 | 5.9543
1216.2 | 0.0012514 |  | -0.0036244 |  | -0.0024923 |  | 0.0063089 |  | -0.0049304 |  | 1531 | 5.9746
1216.1 | 0.0060635 |  | -0.013889 |  | -0.0060697 |  | -0.0012208 |  | -0.0037117 |  | 1530.5 | 5.9577
1216 | 0.0060987 |  | -0.0075613 |  | -0.0012697 |  | -0.003685 |  | -0.0024549 |  | 1530 | 5.9809
1215.9 | -0.0012601 |  | -0.0073259 |  | -0.0073299 |  | 0.0024667 |  | -0.011133 |  | 1529.5 | 5.9762
1215.8 | -0.0012117 |  | -0.0049107 |  | -0.0048831 |  | -0.00996 |  | -0.0073283 |  | 1529 | 5.9684
1215.7 | 0.0012617 |  | 0.0012353 |  | 0.0025217 |  | -0.017056 |  | -0.013848 |  | 1528.5 | 5.9735
1215.6 | -0.0012636 |  | -0.0088514 |  | 0.0012525 |  | -0.015066 |  | 0 |  | 1528 | 5.9757
1215.5 | 0.0012216 |  | -0.0024186 |  | -0.0024187 |  | -0.0036632 |  | -0.014539 |  | 1527.5 | 5.9781
1215.4 | 0.0036788 |  | -0.01273 |  | -0.0061129 |  | -0.0061193 |  | -0.0088756 |  | 1527 | 5.9905
1215.3 | -0.0024694 |  | -0.012166 |  | -0.0087773 |  | 0 |  | -0.0061029 |  | 1526.5 | 5.9805
1215.2 | 0.0012667 |  | -0.0098116 |  | -0.011346 |  | -0.014502 |  | -0.0073585 |  | 1526 | 5.9825
1215.1 | -0.0012159 |  | -0.011157 |  | -0.010963 |  | -0.0050439 |  | -0.0012591 |  | 1525.5 | 5.9711
1215 | 0.0024625 |  | -0.0062177 |  | -0.0012237 |  | -0.0037802 |  | -0.0099895 |  | 1525 | 5.9878
1214.9 | 0.0012426 |  | -0.0060946 |  | -0.0098388 |  | -0.010877 |  | -0.0024208 |  | 1524.5 | 5.9896
1214.8 | -0.0024461 |  | -0.0087862 |  | -0.0088374 |  | -0.0061144 |  | -0.006122 |  | 1524 | 5.9883
1214.7 | -0.011346 |  | -0.0097729 |  | -0.018106 |  | -0.017334 |  | -0.0088055 |  | 1523.5 | 5.9967
1214.6 | 0.0024249 |  | -0.0012368 |  | -0.0076405 |  | -0.0074537 |  | -0.0048412 |  | 1523 | 5.9841
1214.5 | 0 |  | 0.0012307 |  | -0.0085112 |  | -0.012167 |  | -0.0012606 |  | 1522.5 | 5.9937
1214.4 | -0.0012607 |  | 0.0025158 |  | -0.0036563 |  | 0 |  | 0 |  | 1522 | 6.0124
1214.3 | -0.0012213 |  | -0.0036171 |  | -0.015139 |  | -0.010977 |  | -0.0073257 |  | 1521.5 | 6.0003
1214.2 | 0.001226 |  | -0.0050793 |  | -0.0075128 |  | -0.0037698 |  | -0.0073572 |  | 1521 | 5.9942
1214.1 | -0.0049331 |  | 0.0024417 |  | -0.010928 |  | -0.0024455 |  | -0.012343 |  | 1520.5 | 6.0061
1214 | 0.0037305 |  | -0.0024476 |  | -0.0061503 |  | -0.0048983 |  | -0.012461 |  | 1520 | 6.0088
1213.9 | 0.0060929 |  | -0.0050038 |  | -0.016154 |  | -0.0049552 |  | -0.0036503 |  | 1519.5 | 6.0122
1213.8 | 0.0012551 |  | 0 |  | -0.0048899 |  | -0.0074567 |  | 0.0037526 |  | 1519 | 6.0027
1213.7 | -0.0036625 |  | -0.0037174 |  | -0.0012537 |  | -0.0073109 |  | -0.0085267 |  | 1518.5 | 6.0122
1213.6 | -0.0061129 |  | 0 |  | -0.0062979 |  | -0.010031 |  | -0.0062659 |  | 1518 | 6.0092
1213.5 | 0.0012537 |  | -0.0062147 |  | -0.012167 |  | -0.0036618 |  | -0.0012218 |  | 1517.5 | 6.0063
1213.4 | -0.0012075 |  | 0.0048757 |  | -0.0073312 |  | -0.0012363 |  | -0.0073362 |  | 1517 | 6.0212
1213.3 | -0.0025183 |  | -0.0037884 |  | -0.0098525 |  | -0.0024714 |  | -0.0098743 |  | 1516.5 | 6.0218
1213.2 | -0.002515 |  | 0.0036883 |  | -0.0050541 |  | -0.0049445 |  | 0.0037239 |  | 1516 | 6.0126
1213.1 | -0.0012164 |  | -0.0036695 |  | -0.0072228 |  | 0.0012352 |  | -0.0060798 |  | 1515.5 | 6.0043
1213 | -0.0012314 |  | -0.0036752 |  | -0.0089007 |  | 0 |  | 0.002501 |  | 1515 | 6.0216
1212.9 | -0.0049747 |  | -0.0049751 |  | -0.0097324 |  | 0.0024728 |  | -0.0036547 |  | 1514.5 | 6.0245
1212.8 | 0.0024479 |  | -0.0024521 |  | -0.0048828 |  | 0.0037119 |  | 0.0012355 |  | 1514 | 6.0236
1212.7 | 0 |  | 0.002495 |  | -0.0088301 |  | -0.0024746 |  | -0.0061768 |  | 1513.5 | 6.021
1212.6 | 0.0012128 |  | -0.0024583 |  | -0.008763 |  | 0 |  | -0.0012346 |  | 1513 | 6.009
1212.5 | -0.0075384 |  | 0.0099039 |  | -0.0036465 |  | -0.0012373 |  | -0.0024704 |  | 1512.5 | 6.0104
1212.4 | 0.0024431 |  | -0.0061442 |  | -0.014782 |  | -0.0037153 |  | -0.0012365 |  | 1512 | 6.0221
1212.3 | 0.0012472 |  | -0.0086848 |  | 0.0049702 |  | 0.0012317 |  | 0.0024716 |  | 1511.5 | 6.0201
1212.2 | 0.0024544 |  | 0.009763 |  | -0.0061283 |  | 0.0050335 |  | -0.0049431 |  | 1511 | 6.0353
1212.1 | -0.0061806 |  | -0.0088195 |  | 0.0012385 |  | 0.0048354 |  | 0.004947 |  | 1510.5 | 6.0355
1212 | 0.0024657 |  | 0.0024459 |  | -0.011186 |  | -0.0025476 |  | -0.0049448 |  | 1510 | 6.0344
1211.9 | 0.0024823 |  | 0.0012389 |  | -0.0048701 |  | 0 |  | 0.0024727 |  | 1509.5 | 6.0368
1211.8 | 0.0012269 |  | 0.011059 |  | -0.011282 |  | -0.0012295 |  | -0.0037082 |  | 1509 | 6.0271
1211.7 | -0.0012349 |  | 0.0062035 |  | -0.010963 |  | -0.0012417 |  | 0.002471 |  | 1508.5 | 6.0231
1211.6 | 0.0024722 |  | -0.0012203 |  | -0.013597 |  | -0.0036971 |  | -0.0061842 |  | 1508 | 6.0522
1211.5 | 0 |  | -0.0012606 |  | -0.0061787 |  | 0.002471 |  | -0.0098936 |  | 1507.5 | 6.0341
1211.4 | 0.0037078 |  | -0.0024461 |  | -0.0024712 |  | 0.0037473 |  | -0.0012365 |  | 1507 | 6.0521
1211.3 | -0.0074189 |  | 0.0061788 |  | -0.0024711 |  | 0.0060987 |  | 0.0024729 |  | 1506.5 | 6.0322
1211.2 | 0.0049443 |  | 0 |  | -0.0012351 |  | 0.0087687 |  | 0.0037075 |  | 1506 | 6.0293
1211.1 | 0 |  | 0.0012361 |  | -0.0037127 |  | 0.0060574 |  | 0 |  | 1505.5 | 6.0352
1211 | 0.0050134 |  | 0.0024738 |  | 0.0012382 |  | -0.0025385 |  | -0.001258 |  | 1505 | 6.1012
1210.9 | 0.012142 |  | 0.0073924 |  | -0.004945 |  | -0.0012133 |  | 0 |  | 1504.5 | 6.0282
1210.8 | 0.0062909 |  | 0.0074783 |  | -0.0086538 |  | 0.002455 |  | 0.0037803 |  | 1504 | 6.0317
1210.7 | -0.0024324 |  | 0.0024317 |  | -0.0037061 |  | 0.0024892 |  | 0.0024171 |  | 1503.5 | 6.0388
1210.6 | -0.0012494 |  | 0 |  | -0.002473 |  | 0.0085614 |  | -0.0050874 |  | 1503 | 6.0454
1210.5 | 0.00122 |  | 0.0024195 |  | 0.0012364 |  | 0.0037336 |  | 0.0024379 |  | 1502.5 | 6.0552
1210.4 | -0.0050238 |  | -0.0025499 |  | 0.0074181 |  | 0.0061215 |  | 0.0061775 |  | 1502 | 6.0375
1210.3 | 0.001222 |  | 0.004877 |  | 0.0012362 |  | 0.0024902 |  | 0 |  | 1501.5 | 6.0504
1210.2 | 0 |  | 0.0061493 |  | 0.0086503 |  | -0.0012291 |  | 0.0024674 |  | 1501 | 6.0481
1210.1 | 0.0025071 |  | 0.0061909 |  | 0.0074189 |  | 0.0024705 |  | 0.0049738 |  | 1500.5 | 6.0308
1210 | 0.0024185 |  | -0.0024813 |  | 0.0024749 |  | 0.0024726 |  | 0.0049299 |  | 1500 | 6.0417
1209.9 | 0.015114 |  | -0.0086229 |  | 0.0061813 |  | 0.01235 |  | -0.0024863 |  | 1499.5 | 6.0358
1209.8 | -0.011296 |  | 0.004917 |  | 0 |  | 0.0049416 |  | 0.0024398 |  | 1499 | 6.0311
1209.7 | 0.0048334 |  | 0.0037084 |  | 0.0012386 |  | 0.004948 |  | 0 |  | 1498.5 | 6.0473
1209.6 | -0.002448 |  | 0.0037118 |  | -0.0024589 |  | -0.0024756 |  | 0.0012372 |  | 1498 | 6.0405
1209.5 | -0.0025082 |  | 0.017252 |  | 0.0037241 |  | -0.0024747 |  | 0.0012172 |  | 1497.5 | 6.0349
1209.4 | 0.0036251 |  | 0.0012429 |  | 0.0012187 |  | 0.0037084 |  | 0 |  | 1497 | 6.0522
1209.3 | 0 |  | 0.011034 |  | 0.0062991 |  | 0.011124 |  | 0 |  | 1496.5 | 6.0268
1209.2 | 0.0037653 |  | 0.0012273 |  | 0.013441 |  | 0.0061857 |  | 0.0075234 |  | 1496 | 6.0386
1209.1 | 0.0072498 |  | -0.0074503 |  | 0.0098911 |  | 0.007423 |  | 0.0061037 |  | 1495.5 | 6.0309
1209 | 0.0086542 |  | 0.0061655 |  | -0.0037112 |  | 0.0012376 |  | 0.0037156 |  | 1495 | 6.0397
1208.9 | 0.0073872 |  | -0.0012298 |  | 0.0049244 |  | 0.011098 |  | 0.013545 |  | 1494.5 | 6.03
1208.8 | -0.0024882 |  | 0 |  | -0.0024919 |  | 0.0098533 |  | 0.0012293 |  | 1494 | 6.0308
1208.7 | 0.0036496 |  | -0.0049494 |  | 0.0060834 |  | 0.0086353 |  | 0.0086918 |  | 1493.5 | 6.0322
1208.6 | -0.0012593 |  | 0 |  | 0.0025236 |  | 0.0098542 |  | 0.0098735 |  | 1493 | 6.0209
1208.5 | 0.0060378 |  | 0.0024796 |  | 0.0060386 |  | 0 |  | 0.0024778 |  | 1492.5 | 6.0482
1208.4 | 0.0012726 |  | -0.0073828 |  | 0.0063723 |  | -0.0012354 |  | 0.0012278 |  | 1492 | 6.0461
1208.3 | -0.0060944 |  | -0.0074125 |  | -0.0012181 |  | 0 |  | 0.016161 |  | 1491.5 | 6.0348
1208.2 | 0.003689 |  | -0.0012343 |  | 0.0061007 |  | -0.0061519 |  | 0.0036921 |  | 1491 | 6.0469
1208.1 | 0.001238 |  | -0.0024755 |  | 0 |  | 0.0012409 |  | 0 |  | 1490.5 | 6.0327
1208 | 0.011166 |  | 0.0012273 |  | 0 |  | -2.2411e-07 |  | -0.0037329 |  | 1490 | 6.0462
1207.9 | -0.0012331 |  | 0.0037406 |  | 0.0025301 |  | 0.0099572 |  | 0.0036529 |  | 1489.5 | 6.0529
1207.8 | 0 |  | 2.229e-07 |  | 0.001244 |  | 0.0098338 |  | 0.008804 |  | 1489 | 6.0589
1207.7 | 0.0037098 |  | 0.0050173 |  | 0.0012139 |  | -0.0074086 |  | -0.0012205 |  | 1488.5 | 6.0657
1207.6 | 0.0024732 |  | 0.0048458 |  | 0.0038201 |  | 0 |  | 0.0087559 |  | 1488 | 6.0571
1207.5 | 0.0074145 |  | 0.0063222 |  | 0.0084895 |  | 0 |  | 0.0060981 |  | 1487.5 | 6.0474
1207.4 | 0.0049469 |  | 0.0073129 |  | 0.0049195 |  | 0.0037049 |  | 0.0037641 |  | 1487 | 6.0718
1207.3 | 0.0061847 |  | 0.001246 |  | 0.0037272 |  | 0 |  | 0.0036679 |  | 1486.5 | 6.0722
1207.2 | 0.0061912 |  | -0.0036666 |  | 0 |  | 0.0086569 |  | 0.0061421 |  | 1486 | 6.1134
1207.1 | 0.0036874 |  | 2.2779e-07 |  | 0.0012638 |  | 0.0037078 |  | -0.0049803 |  | 1485.5 | 6.0798
1207 | 0.0012282 |  | 0.0024382 |  | 0.0024488 |  | 0.0024719 |  | -0.009765 |  | 1485 | 6.0896
1206.9 | 0 |  | -0.0012569 |  | 0.0049602 |  | 0.0024717 |  | 0 |  | 1484.5 | 6.1274
1206.8 | 0.0037284 |  | 0.0024424 |  | 0.0098584 |  | 0 |  | 0.0048894 |  | 1484 | 6.1285
1206.7 | 0.0024383 |  | -0.0024994 |  | 0.008704 |  | -0.0012366 |  | 0.0050058 |  | 1483.5 | 6.1291
1206.6 | 0.012497 |  | -0.0012234 |  | 0.0012226 |  | 0.0049458 |  | 0.0061176 |  | 1483 | 6.1395
1206.5 | -0.0036683 |  | 0 |  | 0.003792 |  | 0.0086551 |  | 0 |  | 1482.5 | 6.1507
1206.4 | 0.0049372 |  | 0.0036469 |  | -0.0024501 |  | 0.0061826 |  | -0.0012202 |  | 1482 | 6.1734
1206.3 | 0 |  | 0.0037972 |  | 2.243e-07 |  | 0.0012376 |  | 0.0050199 |  | 1481.5 | 6.1962
1206.2 | 0 |  | -0.001212 |  | 0.0086233 |  | -0.0037125 |  | 0.003665 |  | 1481 | 6.199
1206.1 | 0.0024711 |  | 0.0025325 |  | 0.0037081 |  | 0.0086013 |  | 0.0075434 |  | 1480.5 | 6.2352
1206 | -0.0024707 |  | 0.0036615 |  | 0.009876 |  | 0.002459 |  | 0.0036772 |  | 1480 | 6.2234
1205.9 | 0.0061768 |  | 0.0061722 |  | 0.0012389 |  | 0.0074084 |  | 0.0012449 |  | 1479.5 | 6.2022
1205.8 | 0.0037077 |  | 0.0061702 |  | -0.0036891 |  | -0.0036936 |  | -0.0024591 |  | 1479 | 6.2366
1205.7 | 0 |  | 0.0049671 |  | -0.0012417 |  | -0.004951 |  | -0.004945 |  | 1478.5 | 6.2578
1205.6 | 0.0037065 |  | -0.0024649 |  | 0.0086278 |  | 0.0061586 |  | -0.0049457 |  | 1478 | 6.3044
1205.5 | 0 |  | 0.0024836 |  | 0.01612 |  | -0.0037105 |  | -0.0098997 |  | 1477.5 | 6.3146
1205.4 | 0.0012365 |  | 0.0012346 |  | 0.004909 |  | -0.001237 |  | -0.0012368 |  | 1477 | 6.3525
1205.3 | 0.0012368 |  | 0.004948 |  | -0.0012403 |  | -0.0037114 |  | 0.0086603 |  | 1476.5 | 6.365
1205.2 | 0.0037108 |  | 0.0024636 |  | 0.012302 |  | 0.0012366 |  | 0.0012373 |  | 1476 | 6.3795
1205.1 | 0.0024631 |  | -0.009876 |  | 0.0074587 |  | 0.0036957 |  | 0.0098396 |  | 1475.5 | 6.364
1205 | 0.0024617 |  | 0.0012331 |  | 0 |  | 0.0098942 |  | 0.0036875 |  | 1475 | 6.3833
1204.9 | -0.0012346 |  | 0.0024703 |  | 0.0086305 |  | -0.0073514 |  | 0.0012392 |  | 1474.5 | 6.4368
1204.8 | 0.0074254 |  | -0.0024712 |  | 0.0012473 |  | 0.0012439 |  | 0.0037041 |  | 1474 | 6.4494
1204.7 | 0.0036789 |  | 0.0085596 |  | 0.0061004 |  | 0.0073517 |  | 0.0061269 |  | 1473.5 | 6.4984
1204.6 | 0.004901 |  | 0.0037282 |  | 0.014943 |  | 0.0099375 |  | -0.002451 |  | 1473 | 6.4715
1204.5 | 0.0049351 |  | -0.0012262 |  | 0.00122 |  | 0.014708 |  | -0.0012337 |  | 1472.5 | 6.5172
1204.4 | 0.007478 |  | 0.0024653 |  | 0.018973 |  | 0.0012488 |  | 0.011236 |  | 1472 | 6.5119
1204.3 | 0.006109 |  | 0.0037025 |  | 0.01213 |  | 0.0036587 |  | 0.0097843 |  | 1471.5 | 6.512
1204.2 | 0.016121 |  | 0.011107 |  | 0.0050048 |  | 0.0037411 |  | 0.0012443 |  | 1471 | 6.555
1204.1 | 0.0097798 |  | 0.0037044 |  | 0.002433 |  | 0 |  | 0.0085616 |  | 1470.5 | 6.5437
1204 | 0.0087009 |  | 0.0024711 |  | 0.018523 |  | 0.0050156 |  | 0.0087099 |  | 1470 | 6.6165
1203.9 | 0.0036758 |  | 0.0061748 |  | 0.00247 |  | -0.0012252 |  | 0.0098096 |  | 1469.5 | 6.5995
1203.8 | 0.013714 |  | 0.0074065 |  | 0.020993 |  | 0 |  | 0.014703 |  | 1469 | 6.6278
1203.7 | 0.015966 |  | -0.0012355 |  | 0.017314 |  | 0.0073533 |  | 0.0086441 |  | 1468.5 | 6.66
1203.6 | 0.017207 |  | 0.004943 |  | 0.0074325 |  | 0.0074771 |  | 0.016087 |  | 1468 | 6.6308
1203.5 | 0.013651 |  | 0.0012364 |  | 0.0098725 |  | 0.0098438 |  | 0.0073525 |  | 1467.5 | 6.6946
1203.4 | 0.012249 |  | 0.0049499 |  | 0.015947 |  | 0.0049439 |  | 0.009804 |  | 1467 | 6.6967
1203.3 | 0.006164 |  | 0.014718 |  | 0.0074003 |  | 0.0073732 |  | 0.0012374 |  | 1466.5 | 6.6939
1203.2 | 0.015982 |  | 0.0099346 |  | 0.0074096 |  | 0.0036842 |  | 0.012265 |  | 1466 | 6.7125
1203.1 | 0.021025 |  | 0.01103 |  | 0.022088 |  | 0.0061761 |  | 0.011192 |  | 1465.5 | 6.7314
1203 | 0.020938 |  | 0.012352 |  | 0.0085918 |  | 0.0061763 |  | 0.013507 |  | 1465 | 6.8317
1202.9 | 0.023534 |  | 0.0086003 |  | 0.01854 |  | 0.012301 |  | 0.0062143 |  | 1464.5 | 6.7194
1202.8 | 0.020946 |  | 0.0049126 |  | 0.019588 |  | 0.01723 |  | 0.012328 |  | 1464 | 6.8426
1202.7 | 0.014833 |  | 0.012384 |  | 0.01241 |  | 0.0074161 |  | 0.013497 |  | 1463.5 | 6.7488
1202.6 | 0.021035 |  | 0.013496 |  | 0.028259 |  | 0.018501 |  | 0.018426 |  | 1463 | 6.7477
1202.5 | 0.024724 |  | 0.0086834 |  | 0.018634 |  | 0.022261 |  | 0.011116 |  | 1462.5 | 6.8929
1202.4 | 0.02352 |  | 0.0036852 |  | 0.024652 |  | 0.020999 |  | 0.017343 |  | 1462 | 6.7935
1202.3 | 0.024557 |  | 0.01491 |  | 0.028436 |  | 0.012307 |  | 0.013465 |  | 1461.5 | 6.842
1202.2 | 0.030707 |  | 0.019688 |  | 0.023487 |  | 0.019673 |  | 0.019913 |  | 1461 | 6.8575
1202.1 | 0.026066 |  | 0.016044 |  | 0.022264 |  | 0.019722 |  | 0.019611 |  | 1460.5 | 6.8629
1202 | 0.023425 |  | 0.02591 |  | 0.024726 |  | 0.015981 |  | 0.018676 |  | 1460 | 6.8048
1201.9 | 0.030841 |  | 0.018512 |  | 0.029469 |  | 0.018533 |  | 0.017193 |  | 1459.5 | 6.8438
1201.8 | 0.038403 |  | 0.0086433 |  | 0.030714 |  | 0.021035 |  | 0.013508 |  | 1459 | 6.8798
1201.7 | 0.034461 |  | 0.020995 |  | 0.024781 |  | 0.023328 |  | 0.021017 |  | 1458.5 | 6.8784
1201.6 | 0.03463 |  | 0.017279 |  | 0.028262 |  | 0.019893 |  | 0.01858 |  | 1458 | 6.9602
1201.5 | 0.031975 |  | 0.020987 |  | 0.028549 |  | 0.029491 |  | 0.020825 |  | 1457.5 | 6.8867
1201.4 | 0.042982 |  | 0.034597 |  | 0.036818 |  | 0.028321 |  | 0.029792 |  | 1457 | 6.9868
1201.3 | 0.040759 |  | 0.024709 |  | 0.026121 |  | 0.028611 |  | 0.029428 |  | 1456.5 | 6.9304
1201.2 | 0.032141 |  | 0.029538 |  | 0.032001 |  | 0.034404 |  | 0.026219 |  | 1456 | 6.912
1201.1 | 0.046774 |  | 0.022275 |  | 0.037032 |  | 0.035947 |  | 0.029363 |  | 1455.5 | 6.9235
1201 | 0.052984 |  | 0.034342 |  | 0.040744 |  | 0.032171 |  | 0.038238 |  | 1455 | 7.0005
1200.9 | 0.046976 |  | 0.029702 |  | 0.044437 |  | 0.037004 |  | 0.041734 |  | 1454.5 | 6.9864
1200.8 | 0.057992 |  | 0.039363 |  | 0.034978 |  | 0.045592 |  | 0.036829 |  | 1454 | 6.9959
1200.7 | 0.051601 |  | 0.033544 |  | 0.049017 |  | 0.042004 |  | 0.051939 |  | 1453.5 | 6.8577
1200.6 | 0.048597 |  | 0.041741 |  | 0.056842 |  | 0.044372 |  | 0.049535 |  | 1453 | 6.8906
1200.5 | 0.0589 |  | 0.033428 |  | 0.061531 |  | 0.058874 |  | 0.049043 |  | 1452.5 | 7.063
1200.4 | 0.060483 |  | 0.038348 |  | 0.051691 |  | 0.054489 |  | 0.055892 |  | 1452 | 6.8919
1200.3 | 0.062712 |  | 0.039446 |  | 0.050695 |  | 0.052878 |  | 0.05891 |  | 1451.5 | 6.8829
1200.2 | 0.065225 |  | 0.049364 |  | 0.064086 |  | 0.064617 |  | 0.063551 |  | 1451 | 6.9165
1200.1 | 0.072774 |  | 0.050867 |  | 0.066733 |  | 0.061459 |  | 0.058945 |  | 1450.5 | 7.0401
1200 | 0.083668 |  | 0.056827 |  | 0.075219 |  | 0.062887 |  | 0.075394 |  | 1450 | 6.9724
1199.9 | 0.084238 |  | 0.05504 |  | 0.054132 |  | 0.05875 |  | 0.064799 |  | 1449.5 | 7.0765
1199.8 | 0.091632 |  | 0.069099 |  | 0.067835 |  | 0.070202 |  | 0.075384 |  | 1449 | 6.973
1199.7 | 0.086192 |  | 0.072575 |  | 0.068767 |  | 0.07357 |  | 0.078132 |  | 1448.5 | 7.0252
1199.6 | 0.10018 |  | 0.068419 |  | 0.076822 |  | 0.085649 |  | 0.084883 |  | 1448 | 7.0577
1199.5 | 0.10156 |  | 0.069724 |  | 0.083032 |  | 0.086419 |  | 0.095303 |  | 1447.5 | 7.0273
1199.4 | 0.10943 |  | 0.085188 |  | 0.081253 |  | 0.088393 |  | 0.09462 |  | 1447 | 6.966
1199.3 | 0.1193 |  | 0.093481 |  | 0.088702 |  | 0.081691 |  | 0.10565 |  | 1446.5 | 6.9406
1199.2 | 0.13325 |  | 0.09301 |  | 0.1001 |  | 0.10551 |  | 0.11012 |  | 1446 | 6.8912
1199.1 | 0.13472 |  | 0.10002 |  | 0.10749 |  | 0.11415 |  | 0.12193 |  | 1445.5 | 6.9582
1199 | 0.14276 |  | 0.097307 |  | 0.11445 |  | 0.10517 |  | 0.12631 |  | 1445 | 7.0365
1198.9 | 0.15782 |  | 0.096098 |  | 0.12446 |  | 0.11768 |  | 0.12935 |  | 1444.5 | 7.023
1198.8 | 0.17342 |  | 0.11097 |  | 0.12924 |  | 0.12905 |  | 0.13408 |  | 1444 | 7.0138
1198.7 | 0.17374 |  | 0.12331 |  | 0.14159 |  | 0.15065 |  | 0.15009 |  | 1443.5 | 6.9868
1198.6 | 0.18572 |  | 0.12706 |  | 0.1452 |  | 0.15498 |  | 0.16329 |  | 1443 | 6.9445
1198.5 | 0.19715 |  | 0.13702 |  | 0.15292 |  | 0.16145 |  | 0.17475 |  | 1442.5 | 6.8917
1198.4 | 0.21578 |  | 0.14266 |  | 0.17469 |  | 0.17719 |  | 0.18125 |  | 1442 | 6.9624
1198.3 | 0.22419 |  | 0.15624 |  | 0.18687 |  | 0.18183 |  | 0.19952 |  | 1441.5 | 6.8836
1198.2 | 0.25476 |  | 0.16428 |  | 0.19768 |  | 0.1974 |  | 0.20751 |  | 1441 | 6.9623
1198.1 | 0.27229 |  | 0.17353 |  | 0.2125 |  | 0.21231 |  | 0.22108 |  | 1440.5 | 6.9302
1198 | 0.29451 |  | 0.18983 |  | 0.22979 |  | 0.2331 |  | 0.23867 |  | 1440 | 7.0398
1197.9 | 0.31797 |  | 0.19981 |  | 0.24208 |  | 0.25173 |  | 0.26005 |  | 1439.5 | 6.9525
1197.8 | 0.33906 |  | 0.20627 |  | 0.26017 |  | 0.2605 |  | 0.28503 |  | 1439 | 6.8593
1197.7 | 0.3578 |  | 0.23686 |  | 0.28151 |  | 0.28293 |  | 0.31226 |  | 1438.5 | 6.9053
1197.6 | 0.38664 |  | 0.2433 |  | 0.30166 |  | 0.29969 |  | 0.33571 |  | 1438 | 6.9262
1197.5 | 0.42453 |  | 0.26548 |  | 0.32728 |  | 0.33905 |  | 0.34673 |  | 1437.5 | 6.8326
1197.4 | 0.45544 |  | 0.28514 |  | 0.35996 |  | 0.36261 |  | 0.38273 |  | 1437 | 6.854
1197.3 | 0.48343 |  | 0.31098 |  | 0.37885 |  | 0.38487 |  | 0.41029 |  | 1436.5 | 6.7982
1197.2 | 0.5286 |  | 0.32151 |  | 0.40342 |  | 0.40995 |  | 0.43464 |  | 1436 | 6.7976
1197.1 | 0.5772 |  | 0.35172 |  | 0.44759 |  | 0.44315 |  | 0.47121 |  | 1435.5 | 6.8172
1197 | 0.6289 |  | 0.38807 |  | 0.48266 |  | 0.48756 |  | 0.51605 |  | 1435 | 6.789
1196.9 | 0.67759 |  | 0.40391 |  | 0.51892 |  | 0.52527 |  | 0.56305 |  | 1434.5 | 6.7183
1196.8 | 0.73732 |  | 0.44535 |  | 0.56177 |  | 0.57185 |  | 0.61532 |  | 1434 | 6.8213
1196.7 | 0.79791 |  | 0.46892 |  | 0.60739 |  | 0.61293 |  | 0.65274 |  | 1433.5 | 6.785
1196.6 | 0.87888 |  | 0.51699 |  | 0.669 |  | 0.67271 |  | 0.71946 |  | 1433 | 6.768
1196.5 | 0.96401 |  | 0.57052 |  | 0.73035 |  | 0.73463 |  | 0.7855 |  | 1432.5 | 6.7376
1196.4 | 1.0359 |  | 0.60779 |  | 0.79624 |  | 0.80852 |  | 0.85129 |  | 1432 | 6.6825
1196.3 | 1.116 |  | 0.66093 |  | 0.8535 |  | 0.87278 |  | 0.91027 |  | 1431.5 | 6.6738
1196.2 | 1.2319 |  | 0.72251 |  | 0.94042 |  | 0.94953 |  | 1.0161 |  | 1431 | 6.6852
1196.1 | 1.3567 |  | 0.79118 |  | 1.0245 |  | 1.0393 |  | 1.1176 |  | 1430.5 | 6.5942
1196 | 1.4639 |  | 0.86353 |  | 1.1207 |  | 1.1287 |  | 1.2126 |  | 1430 | 6.6089
1195.9 | 1.5872 |  | 0.92443 |  | 1.2023 |  | 1.2255 |  | 1.3029 |  | 1429.5 | 6.5183
1195.8 | 1.7425 |  | 1.0078 |  | 1.3152 |  | 1.3464 |  | 1.4177 |  | 1429 | 6.5269
1195.7 | 1.9262 |  | 1.1111 |  | 1.4598 |  | 1.4782 |  | 1.568 |  | 1428.5 | 6.4619
1195.6 | 2.1064 |  | 1.2136 |  | 1.5832 |  | 1.619 |  | 1.723 |  | 1428 | 6.4686
1195.5 | 2.2984 |  | 1.3113 |  | 1.7203 |  | 1.7493 |  | 1.8729 |  | 1427.5 | 6.4918
1195.4 | 2.5022 |  | 1.4311 |  | 1.8783 |  | 1.9083 |  | 2.0563 |  | 1427 | 6.4528
1195.3 | 2.8266 |  | 1.5973 |  | 2.1047 |  | 2.1364 |  | 2.2968 |  | 1426.5 | 6.3962
1195.2 | 3.094 |  | 1.7384 |  | 2.3089 |  | 2.3505 |  | 2.5364 |  | 1426 | 6.3645
1195.1 | 3.4 |  | 1.8957 |  | 2.5136 |  | 2.5838 |  | 2.777 |  | 1425.5 | 6.2801
1195 | 3.7286 |  | 2.0777 |  | 2.7568 |  | 2.8192 |  | 3.0343 |  | 1425 | 6.2804
1194.9 | 4.1911 |  | 2.3305 |  | 3.0959 |  | 3.1548 |  | 3.3812 |  | 1424.5 | 6.2325
1194.8 | 4.6735 |  | 2.5817 |  | 3.4429 |  | 3.5117 |  | 3.7798 |  | 1424 | 6.2385
1194.7 | 5.1141 |  | 2.821 |  | 3.7679 |  | 3.8644 |  | 4.1838 |  | 1423.5 | 6.1748
1194.6 | 5.6416 |  | 3.0942 |  | 4.1325 |  | 4.2608 |  | 4.5899 |  | 1423 | 6.1281
1194.5 | 6.3193 |  | 3.4492 |  | 4.6265 |  | 4.7232 |  | 5.115 |  | 1422.5 | 6.0969
1194.4 | 7.0419 |  | 3.8447 |  | 5.1582 |  | 5.2513 |  | 5.6962 |  | 1422 | 6.0838
1194.3 | 7.7985 |  | 4.233 |  | 5.7108 |  | 5.8678 |  | 6.3439 |  | 1421.5 | 6.0427
1194.2 | 8.5944 |  | 4.6615 |  | 6.3455 |  | 6.4765 |  | 7.0328 |  | 1421 | 5.9969
1194.1 | 9.4961 |  | 5.1037 |  | 6.9861 |  | 7.152 |  | 7.7532 |  | 1420.5 | 5.965
1194 | 10.7787 |  | 5.7908 |  | 7.9384 |  | 8.0671 |  | 8.758 |  | 1420 | 5.9098
1193.9 | 11.9262 |  | 6.4163 |  | 8.7958 |  | 9.0161 |  | 9.7381 |  | 1419.5 | 5.8984
1193.8 | 13.1426 |  | 7.0574 |  | 9.6766 |  | 9.9672 |  | 10.847 |  | 1419 | 5.8746
1193.7 | 14.5708 |  | 7.8147 |  | 10.7721 |  | 11.0162 |  | 11.974 |  | 1418.5 | 5.8375
1193.6 | 16.5055 |  | 8.9251 |  | 12.2269 |  | 12.4841 |  | 13.5161 |  | 1418 | 5.8134
1193.5 | 18.254 |  | 9.8835 |  | 13.5401 |  | 13.9733 |  | 15.0015 |  | 1417.5 | 5.7793
1193.4 | 20.1206 |  | 10.9441 |  | 15.1151 |  | 15.4437 |  | 16.8386 |  | 1417 | 5.7559
1193.3 | 22.2564 |  | 12.115 |  | 16.6723 |  | 17.1159 |  | 18.6103 |  | 1416.5 | 5.7475
1193.2 | 24.7149 |  | 13.5292 |  | 18.6891 |  | 19.0394 |  | 20.6195 |  | 1416 | 5.7148
1193.1 | 27.5958 |  | 15.3544 |  | 21.0678 |  | 21.5962 |  | 23.3287 |  | 1415.5 | 5.6988
1193 | 30.1874 |  | 16.8894 |  | 23.2612 |  | 23.8317 |  | 25.7588 |  | 1415 | 5.6813
1192.9 | 33.0677 |  | 18.7042 |  | 25.5906 |  | 26.1702 |  | 28.2715 |  | 1414.5 | 5.6705
1192.8 | 36.316 |  | 20.9656 |  | 28.4151 |  | 28.9639 |  | 31.1788 |  | 1414 | 5.6572
1192.7 | 39.818 |  | 23.5404 |  | 31.8024 |  | 32.2824 |  | 34.5641 |  | 1413.5 | 5.6371
1192.6 | 43.365 |  | 26.1202 |  | 34.8681 |  | 35.6277 |  | 38.1603 |  | 1413 | 5.6335
1192.5 | 46.7376 |  | 28.568 |  | 38.0197 |  | 38.6498 |  | 41.4546 |  | 1412.5 | 5.6259
1192.4 | 49.9307 |  | 31.1677 |  | 41.067 |  | 41.6252 |  | 44.5372 |  | 1412 | 5.6128
1192.3 | 54.3051 |  | 34.8825 |  | 45.3521 |  | 45.8229 |  | 48.9106 |  | 1411.5 | 5.6093
1192.2 | 57.8928 |  | 38.2884 |  | 49.2499 |  | 49.6439 |  | 52.8678 |  | 1411 | 5.6021
1192.1 | 61.1884 |  | 41.3846 |  | 52.5706 |  | 52.9582 |  | 56.3601 |  | 1410.5 | 5.5986
1192 | 64.7352 |  | 44.7962 |  | 56.2649 |  | 56.4543 |  | 59.8023 |  | 1410 | 5.5998
1191.9 | 68.3798 |  | 48.9785 |  | 60.4341 |  | 60.3897 |  | 63.7305 |  | 1409.5 | 5.5976
1191.8 | 71.6 |  | 53.0049 |  | 64.1935 |  | 64.0435 |  | 67.3916 |  | 1409 | 5.587
1191.7 | 74.6125 |  | 56.3704 |  | 67.4465 |  | 67.4131 |  | 70.8857 |  | 1408.5 | 5.5859
1191.6 | 77.2696 |  | 60.0233 |  | 70.6537 |  | 70.5308 |  | 73.6866 |  | 1408 | 5.5776
1191.5 | 79.6898 |  | 63.4699 |  | 73.7215 |  | 73.4349 |  | 76.4707 |  | 1407.5 | 5.5768
1191.4 | 82.6032 |  | 67.6742 |  | 77.4269 |  | 76.8418 |  | 79.6416 |  | 1407 | 5.5697
1191.3 | 84.4921 |  | 70.656 |  | 79.6926 |  | 79.3637 |  | 82.1277 |  | 1406.5 | 5.5635
1191.2 | 86.2889 |  | 73.8419 |  | 82.2021 |  | 81.6112 |  | 84.1837 |  | 1406 | 5.5606
1191.1 | 87.8238 |  | 76.7116 |  | 84.4417 |  | 83.6569 |  | 85.9972 |  | 1405.5 | 5.562
1191 | 89.4995 |  | 79.8112 |  | 86.554 |  | 85.7787 |  | 87.8718 |  | 1405 | 5.5487
1190.9 | 90.7224 |  | 82.1141 |  | 88.2034 |  | 87.4433 |  | 89.4177 |  | 1404.5 | 5.5384
1190.8 | 91.6548 |  | 84.2087 |  | 89.5353 |  | 88.8486 |  | 90.5808 |  | 1404 | 5.5335
1190.7 | 92.4734 |  | 86.0131 |  | 90.7457 |  | 89.996 |  | 91.4947 |  | 1403.5 | 5.5263
1190.6 | 93.5833 |  | 87.8457 |  | 91.8592 |  | 91.166 |  | 92.4609 |  | 1403 | 5.5122
1190.5 | 94.1628 |  | 89.2889 |  | 92.8847 |  | 92.1264 |  | 93.4105 |  | 1402.5 | 5.5043
1190.4 | 94.6314 |  | 90.4618 |  | 93.7388 |  | 93.014 |  | 93.9476 |  | 1402 | 5.4952
1190.3 | 95.0382 |  | 91.5398 |  | 94.2867 |  | 93.7223 |  | 94.4482 |  | 1401.5 | 5.4883
1190.2 | 95.4759 |  | 92.4736 |  | 94.6982 |  | 94.2335 |  | 94.9481 |  | 1401 | 5.4778
1190.1 | 95.5442 |  | 93.3895 |  | 95.1372 |  | 94.6175 |  | 95.2031 |  | 1400.5 | 5.4793
1190 | 95.8198 |  | 93.9843 |  | 95.3324 |  | 94.9987 |  | 95.3258 |  | 1400 | 5.4697
1189.9 | 96.0064 |  | 94.5401 |  | 95.6457 |  | 95.1796 |  | 95.6875 |  | 1399.5 | 5.4681
1189.8 | 96.0511 |  | 94.921 |  | 95.7689 |  | 95.4973 |  | 95.8004 |  | 1399 | 5.4656
1189.7 | 96.1586 |  | 95.147 |  | 95.9673 |  | 95.6734 |  | 95.9025 |  | 1398.5 | 5.4614
1189.6 | 96.3149 |  | 95.4346 |  | 96.1133 |  | 95.7534 |  | 96.0041 |  | 1398 | 5.4623
1189.5 | 96.2541 |  | 95.7223 |  | 96.1083 |  | 95.8576 |  | 96.0747 |  | 1397.5 | 5.4666
1189.4 | 96.3583 |  | 95.8482 |  | 96.2603 |  | 95.8809 |  | 96.1816 |  | 1397 | 5.4764
1189.3 | 96.3165 |  | 95.8929 |  | 96.2014 |  | 96.0648 |  | 96.183 |  | 1396.5 | 5.4814
1189.2 | 96.4286 |  | 96.0455 |  | 96.2637 |  | 96.0921 |  | 96.1993 |  | 1396 | 5.4965
1189.1 | 96.4316 |  | 96.0167 |  | 96.2938 |  | 96.1963 |  | 96.2716 |  | 1395.5 | 5.5098
1189 | 96.497 |  | 96.159 |  | 96.3326 |  | 96.1594 |  | 96.2786 |  | 1395 | 5.5282
1188.9 | 96.3663 |  | 96.2112 |  | 96.291 |  | 96.1647 |  | 96.3229 |  | 1394.5 | 5.5365
1188.8 | 96.4871 |  | 96.2962 |  | 96.3271 |  | 96.2105 |  | 96.347 |  | 1394 | 5.5551
1188.7 | 96.4157 |  | 96.247 |  | 96.3527 |  | 96.22 |  | 96.396 |  | 1393.5 | 5.5729
1188.6 | 96.4622 |  | 96.2329 |  | 96.3449 |  | 96.1728 |  | 96.378 |  | 1393 | 5.5974
1188.5 | 96.4141 |  | 96.3364 |  | 96.3327 |  | 96.2843 |  | 96.4021 |  | 1392.5 | 5.6238
1188.4 | 96.4507 |  | 96.4171 |  | 96.3758 |  | 96.3476 |  | 96.4921 |  | 1392 | 5.6511
1188.3 | 96.4578 |  | 96.3926 |  | 96.3724 |  | 96.2768 |  | 96.5422 |  | 1391.5 | 5.6716
1188.2 | 96.445 |  | 96.3078 |  | 96.3074 |  | 96.3652 |  | 96.4857 |  | 1391 | 5.6986
1188.1 | 96.4405 |  | 96.4078 |  | 96.399 |  | 96.2353 |  | 96.5027 |  | 1390.5 | 5.7168
1188 | 96.413 |  | 96.5177 |  | 96.3228 |  | 96.3131 |  | 96.4492 |  | 1390 | 5.748
1187.9 | 96.4323 |  | 96.4059 |  | 96.4138 |  | 96.2963 |  | 96.5335 |  | 1389.5 | 5.7783
1187.8 | 96.4303 |  | 96.4186 |  | 96.356 |  | 96.2882 |  | 96.6274 |  | 1389 | 5.796
1187.7 | 96.368 |  | 96.3656 |  | 96.3658 |  | 96.3199 |  | 96.4529 |  | 1388.5 | 5.8163
1187.6 | 96.4017 |  | 96.5802 |  | 96.3453 |  | 96.286 |  | 96.5974 |  | 1388 | 5.8425
1187.5 | 96.4795 |  | 96.4735 |  | 96.4567 |  | 96.3845 |  | 96.6372 |  | 1387.5 | 5.8557
1187.4 | 96.555 |  | 96.3668 |  | 96.3536 |  | 96.3432 |  | 96.6071 |  | 1387 | 5.8789
1187.3 | 96.4116 |  | 96.5447 |  | 96.2864 |  | 96.38 |  | 96.5941 |  | 1386.5 | 5.865
1187.2 | 96.4059 |  | 96.598 |  | 96.4109 |  | 96.3696 |  | 96.5761 |  | 1386 | 5.8849
1187.1 | 96.4652 |  | 96.6488 |  | 96.3055 |  | 96.3776 |  | 96.6307 |  | 1385.5 | 5.8775
1187 | 96.4559 |  | 96.5875 |  | 96.3905 |  | 96.3557 |  | 96.6326 |  | 1385 | 5.8799
1186.9 | 96.4332 |  | 96.4604 |  | 96.3178 |  | 96.3266 |  | 96.5566 |  | 1384.5 | 5.8704
1186.8 | 96.4952 |  | 96.5725 |  | 96.3075 |  | 96.2657 |  | 96.5897 |  | 1384 | 5.8407
1186.7 | 96.4648 |  | 96.6574 |  | 96.3499 |  | 96.3589 |  | 96.6074 |  | 1383.5 | 5.8159
1186.6 | 96.3734 |  | 96.558 |  | 96.3239 |  | 96.3221 |  | 96.5152 |  | 1383 | 5.7655
1186.5 | 96.4231 |  | 96.523 |  | 96.3208 |  | 96.339 |  | 96.6053 |  | 1382.5 | 5.7288
1186.4 | 96.2575 |  | 96.436 |  | 96.1849 |  | 96.2005 |  | 96.584 |  | 1382 | 5.6877
1186.3 | 96.3365 |  | 96.5241 |  | 96.2217 |  | 96.181 |  | 96.4971 |  | 1381.5 | 5.6393
1186.2 | 96.494 |  | 96.5797 |  | 96.272 |  | 96.2758 |  | 96.6282 |  | 1381 | 5.6031
1186.1 | 96.4185 |  | 96.5118 |  | 96.2417 |  | 96.2622 |  | 96.5318 |  | 1380.5 | 5.5644
1186 | 96.2811 |  | 96.4887 |  | 96.2492 |  | 96.372 |  | 96.5591 |  | 1380 | 5.528
1185.9 | 96.3536 |  | 96.5397 |  | 96.2619 |  | 96.185 |  | 96.5491 |  | 1379.5 | 5.5006
1185.8 | 96.3596 |  | 96.523 |  | 96.2568 |  | 96.3519 |  | 96.5505 |  | 1379 | 5.4713
1185.7 | 96.3462 |  | 96.4729 |  | 96.2452 |  | 96.2793 |  | 96.587 |  | 1378.5 | 5.4527
1185.6 | 96.3326 |  | 96.5206 |  | 96.2184 |  | 96.2697 |  | 96.5006 |  | 1378 | 5.428
1185.5 | 96.1867 |  | 96.5139 |  | 96.2124 |  | 96.1588 |  | 96.5199 |  | 1377.5 | 5.4175
1185.4 | 96.2214 |  | 96.5416 |  | 96.2538 |  | 96.196 |  | 96.5469 |  | 1377 | 5.4073
1185.3 | 96.152 |  | 96.6325 |  | 96.2532 |  | 96.282 |  | 96.5328 |  | 1376.5 | 5.3954
1185.2 | 96.1975 |  | 96.3803 |  | 96.2164 |  | 96.2069 |  | 96.5229 |  | 1376 | 5.397
1185.1 | 96.2126 |  | 96.4733 |  | 96.242 |  | 96.3067 |  | 96.5543 |  | 1375.5 | 5.396
1185 | 96.143 |  | 96.4783 |  | 96.1786 |  | 96.2343 |  | 96.4961 |  | 1375 | 5.4105
1184.9 | 96.1398 |  | 96.513 |  | 96.2446 |  | 96.1976 |  | 96.5472 |  | 1374.5 | 5.4282
1184.8 | 96.1267 |  | 96.3985 |  | 96.1709 |  | 96.1794 |  | 96.408 |  | 1374 | 5.4495
1184.7 | 96.0219 |  | 96.4416 |  | 96.2323 |  | 96.2389 |  | 96.4656 |  | 1373.5 | 5.4899
1184.6 | 96.1407 |  | 96.4743 |  | 96.304 |  | 96.2449 |  | 96.5631 |  | 1373 | 5.5207
1184.5 | 96.1711 |  | 96.5632 |  | 96.3137 |  | 96.2665 |  | 96.5469 |  | 1372.5 | 5.5658
1184.4 | 96.134 |  | 96.525 |  | 96.2805 |  | 96.3045 |  | 96.545 |  | 1372 | 5.6153
1184.3 | 96.1799 |  | 96.4393 |  | 96.2777 |  | 96.3239 |  | 96.5047 |  | 1371.5 | 5.6607
1184.2 | 96.056 |  | 96.3208 |  | 96.2318 |  | 96.2478 |  | 96.4183 |  | 1371 | 5.7219
1184.1 | 96.2899 |  | 96.4761 |  | 96.3212 |  | 96.3304 |  | 96.4969 |  | 1370.5 | 5.7914
1184 | 96.2174 |  | 96.4904 |  | 96.2762 |  | 96.1398 |  | 96.3513 |  | 1370 | 5.8699
1183.9 | 96.0967 |  | 96.5446 |  | 96.2636 |  | 96.2665 |  | 96.3454 |  | 1369.5 | 5.9499
1183.8 | 96.144 |  | 96.3582 |  | 96.2514 |  | 96.125 |  | 96.3138 |  | 1369 | 6.0311
1183.7 | 96.0867 |  | 96.3429 |  | 96.1481 |  | 96.0608 |  | 96.1561 |  | 1368.5 | 6.1406
1183.6 | 95.9726 |  | 96.255 |  | 96.133 |  | 95.9895 |  | 96.0374 |  | 1368 | 6.2678
1183.5 | 95.9322 |  | 96.2595 |  | 95.9369 |  | 95.8252 |  | 95.8695 |  | 1367.5 | 6.3758
1183.4 | 95.6837 |  | 96.2645 |  | 95.7634 |  | 95.6282 |  | 95.5166 |  | 1367 | 6.4977
1183.3 | 95.4403 |  | 96.1475 |  | 95.5831 |  | 95.4316 |  | 95.3363 |  | 1366.5 | 6.5653
1183.2 | 94.8697 |  | 96.0226 |  | 95.2981 |  | 95.0638 |  | 94.8105 |  | 1366 | 6.6993
1183.1 | 94.404 |  | 95.8459 |  | 94.8138 |  | 94.5538 |  | 94.3274 |  | 1365.5 | 6.9096
1183 | 93.8173 |  | 95.4596 |  | 94.3378 |  | 93.9494 |  | 93.728 |  | 1365 | 6.8494
1182.9 | 92.9967 |  | 95.1875 |  | 93.7195 |  | 93.356 |  | 92.9883 |  | 1364.5 | 6.875
1182.8 | 91.6512 |  | 94.7307 |  | 92.8186 |  | 92.4347 |  | 91.8707 |  | 1364 | 7.0888
1182.7 | 90.2346 |  | 94.2618 |  | 91.9396 |  | 91.3572 |  | 90.4507 |  | 1363.5 | 7.1988
1182.6 | 88.6796 |  | 93.7201 |  | 90.5844 |  | 89.9871 |  | 89.1841 |  | 1363 | 7.164
1182.5 | 86.9621 |  | 92.955 |  | 89.2763 |  | 88.6116 |  | 87.7183 |  | 1362.5 | 7.1672
1182.4 | 84.6493 |  | 91.7262 |  | 87.4905 |  | 86.7637 |  | 85.7503 |  | 1362 | 7.2178
1182.3 | 82.0072 |  | 90.3975 |  | 85.4931 |  | 84.709 |  | 83.5511 |  | 1361.5 | 7.3184
1182.2 | 79.4963 |  | 88.9804 |  | 83.3133 |  | 82.4719 |  | 81.2505 |  | 1361 | 7.0693
1182.1 | 76.5454 |  | 87.4727 |  | 81.1519 |  | 80.0975 |  | 78.6665 |  | 1360.5 | 7.1398
1182 | 73.4957 |  | 85.7439 |  | 78.7277 |  | 77.6869 |  | 76.1227 |  | 1360 | 7.2499
1181.9 | 69.0663 |  | 83.0062 |  | 74.9792 |  | 74.0929 |  | 72.1018 |  | 1359.5 | 7.1807
1181.8 | 65.4182 |  | 80.73 |  | 71.8507 |  | 70.8713 |  | 68.5545 |  | 1359 | 7.1122
1181.7 | 62.0594 |  | 78.3852 |  | 68.8456 |  | 67.8659 |  | 65.7912 |  | 1358.5 | 7.3411
1181.6 | 58.2898 |  | 75.5151 |  | 65.5714 |  | 64.5736 |  | 62.0659 |  | 1358 | 7.1481
1181.5 | 53.1269 |  | 71.5941 |  | 60.9957 |  | 60.148 |  | 57.4493 |  | 1357.5 | 7.4132
1181.4 | 49.628 |  | 68.4898 |  | 57.2559 |  | 56.5407 |  | 53.8491 |  | 1357 | 7.4952
1181.3 | 46.1354 |  | 65.4306 |  | 53.7555 |  | 52.9812 |  | 50.14 |  | 1356.5 | 7.2844
1181.2 | 42.5858 |  | 61.8452 |  | 50.217 |  | 49.6561 |  | 46.732 |  | 1356 | 7.2026
1181.1 | 38.639 |  | 57.4489 |  | 46.386 |  | 46.0669 |  | 43.1413 |  | 1355.5 | 7.3991
1181 | 34.7919 |  | 53.4385 |  | 42.2245 |  | 41.7659 |  | 39.1356 |  | 1355 | 7.1689
1180.9 | 32.0142 |  | 50.3258 |  | 39.194 |  | 38.816 |  | 36.0354 |  | 1354.5 | 7.2787
1180.8 | 29.2006 |  | 46.8408 |  | 36.1531 |  | 35.884 |  | 33.0344 |  | 1354 | 7.5836
1180.7 | 26.4868 |  | 43.1971 |  | 33.1899 |  | 32.9938 |  | 30.2875 |  | 1353.5 | 7.3925
1180.6 | 23.482 |  | 38.9941 |  | 29.4776 |  | 29.5036 |  | 27.0321 |  | 1353 | 7.3974
1180.5 | 21.0524 |  | 35.8635 |  | 26.7559 |  | 26.767 |  | 24.3308 |  | 1352.5 | 7.4395
1180.4 | 19.1252 |  | 32.964 |  | 24.4265 |  | 24.4356 |  | 22.1535 |  | 1352 | 7.6761
1180.3 | 17.3964 |  | 30.2604 |  | 22.3365 |  | 22.4015 |  | 20.297 |  | 1351.5 | 7.3887
1180.2 | 15.1213 |  | 26.9348 |  | 19.6239 |  | 19.8699 |  | 17.8343 |  | 1351 | 7.4574
1180.1 | 13.5464 |  | 24.3505 |  | 17.5922 |  | 17.8321 |  | 15.9405 |  | 1350.5 | 7.4414
1180 | 12.2092 |  | 22.2593 |  | 16.0337 |  | 16.1835 |  | 14.4257 |  | 1350 | 7.6928
1179.9 | 10.9703 |  | 20.1577 |  | 14.5671 |  | 14.663 |  | 13.055 |  | 1349.5 | 7.3953
1179.8 | 9.6492 |  | 17.7439 |  | 12.7707 |  | 13.0186 |  | 11.5431 |  | 1349 | 7.3435
1179.7 | 8.5593 |  | 15.9122 |  | 11.337 |  | 11.6092 |  | 10.2383 |  | 1348.5 | 7.2335
1179.6 | 7.6989 |  | 14.4115 |  | 10.2863 |  | 10.503 |  | 9.2738 |  | 1348 | 7.1089
1179.5 | 6.9855 |  | 13.0266 |  | 9.3081 |  | 9.4918 |  | 8.3569 |  | 1347.5 | 7.3834
1179.4 | 6.1884 |  | 11.644 |  | 8.2875 |  | 8.5275 |  | 7.4782 |  | 1347 | 7.2921
1179.3 | 5.4619 |  | 10.2707 |  | 7.2968 |  | 7.4966 |  | 6.6206 |  | 1346.5 | 7.3457
1179.2 | 4.9212 |  | 9.2652 |  | 6.5943 |  | 6.7524 |  | 5.9358 |  | 1346 | 7.3017
1179.1 | 4.4774 |  | 8.3844 |  | 6.031 |  | 6.1076 |  | 5.4007 |  | 1345.5 | 7.2973
1179 | 4.0616 |  | 7.5562 |  | 5.4454 |  | 5.564 |  | 4.9031 |  | 1345 | 7.7902
1178.9 | 3.5983 |  | 6.7135 |  | 4.8392 |  | 4.9494 |  | 4.3891 |  | 1344.5 | 7.5886
1178.8 | 3.2348 |  | 6.0334 |  | 4.3404 |  | 4.4339 |  | 3.9144 |  | 1344 | 7.0437
1178.7 | 2.9225 |  | 5.4313 |  | 3.9262 |  | 4.024 |  | 3.5317 |  | 1343.5 | 7.3211
1178.6 | 2.6799 |  | 4.9275 |  | 3.5799 |  | 3.6744 |  | 3.2486 |  | 1343 | 7.2873
1178.5 | 2.3908 |  | 4.3879 |  | 3.1928 |  | 3.2724 |  | 2.9223 |  | 1342.5 | 7.4619
1178.4 | 2.1231 |  | 3.9089 |  | 2.8531 |  | 2.9197 |  | 2.6005 |  | 1342 | 7.0869
1178.3 | 1.9518 |  | 3.5668 |  | 2.5866 |  | 2.6648 |  | 2.3719 |  | 1341.5 | 7.2985
1178.2 | 1.7817 |  | 3.2351 |  | 2.3528 |  | 2.4029 |  | 2.1636 |  | 1341 | 7.0256
1178.1 | 1.6186 |  | 2.9108 |  | 2.1251 |  | 2.1837 |  | 1.9591 |  | 1340.5 | 7.2395
1178 | 1.4333 |  | 2.5814 |  | 1.8989 |  | 1.9353 |  | 1.746 |  | 1340 | 7.0693
1177.9 | 1.3031 |  | 2.3332 |  | 1.7276 |  | 1.7549 |  | 1.5812 |  | 1339.5 | 7.1627
1177.8 | 1.2059 |  | 2.1526 |  | 1.5822 |  | 1.6118 |  | 1.4503 |  | 1339 | 7.2561
1177.7 | 1.0944 |  | 1.945 |  | 1.4466 |  | 1.4766 |  | 1.3341 |  | 1338.5 | 7.2505
1177.6 | 0.99218 |  | 1.7443 |  | 1.2953 |  | 1.3171 |  | 1.1948 |  | 1338 | 7.2719
1177.5 | 0.90826 |  | 1.5886 |  | 1.1854 |  | 1.2189 |  | 1.0868 |  | 1337.5 | 7.1291
1177.4 | 0.83563 |  | 1.4521 |  | 1.0706 |  | 1.1141 |  | 1.0008 |  | 1337 | 7.2478
1177.3 | 0.77653 |  | 1.3374 |  | 1.013 |  | 1.0277 |  | 0.91565 |  | 1336.5 | 7.1624
1177.2 | 0.70429 |  | 1.1973 |  | 0.90728 |  | 0.93949 |  | 0.83322 |  | 1336 | 7.0371
1177.1 | 0.63726 |  | 1.0783 |  | 0.81951 |  | 0.85359 |  | 0.75047 |  | 1335.5 | 7.2393
1177 | 0.59967 |  | 0.99744 |  | 0.75703 |  | 0.77927 |  | 0.70838 |  | 1335 | 7.0508
1176.9 | 0.54734 |  | 0.92315 |  | 0.70317 |  | 0.71897 |  | 0.66016 |  | 1334.5 | 7.0578
1176.8 | 0.50378 |  | 0.83817 |  | 0.63465 |  | 0.65454 |  | 0.58551 |  | 1334 | 7.03
1176.7 | 0.46661 |  | 0.75886 |  | 0.58449 |  | 0.59029 |  | 0.54215 |  | 1333.5 | 7.3193
1176.6 | 0.42368 |  | 0.68728 |  | 0.52893 |  | 0.54725 |  | 0.49734 |  | 1333 | 7.1338
1176.5 | 0.39191 |  | 0.63988 |  | 0.49344 |  | 0.50934 |  | 0.46094 |  | 1332.5 | 7.0486
1176.4 | 0.3582 |  | 0.59042 |  | 0.45841 |  | 0.46631 |  | 0.42669 |  | 1332 | 7.1422
1176.3 | 0.32372 |  | 0.52857 |  | 0.4144 |  | 0.42929 |  | 0.38345 |  | 1331.5 | 6.9782
1176.2 | 0.30286 |  | 0.48407 |  | 0.37477 |  | 0.38872 |  | 0.33809 |  | 1331 | 6.9414
1176.1 | 0.28692 |  | 0.44581 |  | 0.35436 |  | 0.36924 |  | 0.33077 |  | 1330.5 | 6.9667
1176 | 0.26604 |  | 0.4245 |  | 0.324 |  | 0.3423 |  | 0.30384 |  | 1330 | 6.8552
1175.9 | 0.23523 |  | 0.37993 |  | 0.30414 |  | 0.31139 |  | 0.2804 |  | 1329.5 | 6.9349
1175.8 | 0.22162 |  | 0.3491 |  | 0.28067 |  | 0.27855 |  | 0.2532 |  | 1329 | 6.9381
1175.7 | 0.20825 |  | 0.32108 |  | 0.25754 |  | 0.26844 |  | 0.24729 |  | 1328.5 | 6.9443
1175.6 | 0.19101 |  | 0.30876 |  | 0.23905 |  | 0.24483 |  | 0.22149 |  | 1328 | 7.0704
1175.5 | 0.17742 |  | 0.27546 |  | 0.22056 |  | 0.22228 |  | 0.20665 |  | 1327.5 | 6.8487
1175.4 | 0.16013 |  | 0.25147 |  | 0.20767 |  | 0.2027 |  | 0.19197 |  | 1327 | 6.9501
1175.3 | 0.15147 |  | 0.2343 |  | 0.19614 |  | 0.19299 |  | 0.1751 |  | 1326.5 | 6.9081
1175.2 | 0.13392 |  | 0.21795 |  | 0.18744 |  | 0.18342 |  | 0.16466 |  | 1326 | 6.7813
1175.1 | 0.13907 |  | 0.19725 |  | 0.15937 |  | 0.16788 |  | 0.15245 |  | 1325.5 | 6.8069
1175 | 0.12834 |  | 0.18516 |  | 0.14763 |  | 0.15558 |  | 0.14484 |  | 1325 | 6.835
1174.9 | 0.11833 |  | 0.18003 |  | 0.1381 |  | 0.1413 |  | 0.12884 |  | 1324.5 | 6.7546
1174.8 | 0.10886 |  | 0.16817 |  | 0.13795 |  | 0.12901 |  | 0.119 |  | 1324 | 6.7785
1174.7 | 0.087668 |  | 0.15144 |  | 0.12342 |  | 0.14007 |  | 0.11827 |  | 1323.5 | 6.9079
1174.6 | 0.078245 |  | 0.13157 |  | 0.11841 |  | 0.11186 |  | 0.10045 |  | 1323 | 6.7282
1174.5 | 0.076251 |  | 0.13023 |  | 0.10942 |  | 0.10568 |  | 0.092031 |  | 1322.5 | 6.7399
1174.4 | 0.07137 |  | 0.12786 |  | 0.10211 |  | 0.089778 |  | 0.084969 |  | 1322 | 6.737
1174.3 | 0.073826 |  | 0.10579 |  | 0.094731 |  | 0.090997 |  | 0.089643 |  | 1321.5 | 6.6763
1174.2 | 0.074995 |  | 0.099564 |  | 0.093469 |  | 0.086046 |  | 0.072294 |  | 1321 | 6.6642
1174.1 | 0.071377 |  | 0.095889 |  | 0.081248 |  | 0.079977 |  | 0.075073 |  | 1320.5 | 6.6516
1174 | 0.061427 |  | 0.089702 |  | 0.07709 |  | 0.067645 |  | 0.071062 |  | 1320 | 6.659
1173.9 | 0.052671 |  | 0.081147 |  | 0.066132 |  | 0.068688 |  | 0.060185 |  | 1319.5 | 6.5485
1173.8 | 0.05548 |  | 0.076302 |  | 0.073736 |  | 0.068962 |  | 0.062657 |  | 1319 | 6.6013
1173.7 | 0.0442 |  | 0.077563 |  | 0.060121 |  | 0.055346 |  | 0.04795 |  | 1318.5 | 6.5799
1173.6 | 0.04551 |  | 0.069792 |  | 0.051516 |  | 0.055172 |  | 0.061205 |  | 1318 | 6.629
1173.5 | 0.049133 |  | 0.057869 |  | 0.055117 |  | 0.052664 |  | 0.045376 |  | 1317.5 | 6.587
1173.4 | 0.04436 |  | 0.067435 |  | 0.04557 |  | 0.044264 |  | 0.048076 |  | 1317 | 6.5653
1173.3 | 0.034344 |  | 0.072791 |  | 0.051437 |  | 0.051622 |  | 0.047895 |  | 1316.5 | 6.537
1173.2 | 0.024599 |  | 0.057561 |  | 0.047947 |  | 0.042968 |  | 0.048842 |  | 1316 | 6.5098
1173.1 | 0.038057 |  | 0.046651 |  | 0.034275 |  | 0.035617 |  | 0.031909 |  | 1315.5 | 6.4535
1173 | 0.031959 |  | 0.049283 |  | 0.050534 |  | 0.035713 |  | 0.035638 |  | 1315 | 6.4992
1172.9 | 0.023371 |  | 0.03804 |  | 0.031874 |  | 0.036894 |  | 0.040574 |  | 1314.5 | 6.4211
1172.8 | 0.031814 |  | 0.036921 |  | 0.039416 |  | 0.029356 |  | 0.035657 |  | 1314 | 6.4446
1172.7 | 0.033264 |  | 0.042962 |  | 0.024648 |  | 0.029519 |  | 0.027055 |  | 1313.5 | 6.3991
1172.6 | 0.027107 |  | 0.04309 |  | 0.02812 |  | 0.034294 |  | 0.02706 |  | 1313 | 6.4103
1172.5 | 0.026954 |  | 0.04434 |  | 0.035756 |  | 0.028336 |  | 0.023378 |  | 1312.5 | 6.3517
1172.4 | 0.018406 |  | 0.036924 |  | 0.020775 |  | 0.018382 |  | 0.025819 |  | 1312 | 6.3199
1172.3 | 0.015918 |  | 0.023398 |  | 0.028129 |  | 0.022093 |  | 0.015987 |  | 1311.5 | 6.3304
1172.2 | 0.030691 |  | 0.029351 |  | 0.022156 |  | 0.023345 |  | 0.025853 |  | 1311 | 6.2863
1172.1 | 0.025917 |  | 0.029496 |  | 0.032046 |  | 0.025829 |  | 0.024647 |  | 1310.5 | 6.3069
1172 | 0.014713 |  | 0.024625 |  | 0.026864 |  | 0.019659 |  | 0.02206 |  | 1310 | 6.2763
1171.9 | 0.01228 |  | 0.024633 |  | 0.023467 |  | 0.013507 |  | 0.019619 |  | 1309.5 | 6.2547
1171.8 | 0.018476 |  | 0.023369 |  | 0.020769 |  | 0.022099 |  | 0.011038 |  | 1309 | 6.2295
1171.7 | 0.0073707 |  | 0.024536 |  | 0.020912 |  | 0.013511 |  | 0.022174 |  | 1308.5 | 6.1947
1171.6 | 0.019621 |  | 0.027012 |  | 0.014715 |  | 0.012281 |  | 0.015938 |  | 1308 | 6.1825
1171.5 | 0.017241 |  | 0.017173 |  | 0.022056 |  | 0.020876 |  | 0.020904 |  | 1307.5 | 6.1555
1171.4 | 0.0036929 |  | 0.018446 |  | 0.017211 |  | 0.0135 |  | 0.019662 |  | 1307 | 6.1567
1171.3 | 0.0098474 |  | 0.027004 |  | 0.019668 |  | 0.012274 |  | 0.014747 |  | 1306.5 | 6.1165
1171.2 | 0.0197 |  | 0.01967 |  | 0.015983 |  | 0.014728 |  | 0.013524 |  | 1306 | 6.1017
1171.1 | 0.0098488 |  | 0.017198 |  | 0.017213 |  | 0.0085917 |  | 0.01967 |  | 1305.5 | 6.0906
1171 | 0.014774 |  | 0.027019 |  | 0.0098369 |  | 0.011046 |  | 0.011065 |  | 1305 | 6.0689
1170.9 | 0.016015 |  | 0.022106 |  | 0.019685 |  | 0.019654 |  | 0.018453 |  | 1304.5 | 6.0372
1170.8 | 0.012243 |  | 0.014735 |  | 0.017189 |  | 0.0098363 |  | 0.013487 |  | 1304 | 6.0223
1170.7 | 0.014747 |  | 0.011065 |  | 0.012247 |  | 0.018453 |  | 0.0049242 |  | 1303.5 | 6.0058
1170.6 | 0.011063 |  | 0.018429 |  | 0.0098415 |  | 0.012289 |  | 0.0097987 |  | 1303 | 5.9818
1170.5 | 0.0061328 |  | 0.015991 |  | 0.022042 |  | 0.015977 |  | 0.020925 |  | 1302.5 | 5.9618
1170.4 | 0.011058 |  | 0.017233 |  | 0.014795 |  | 0.0024592 |  | 0.0098105 |  | 1302 | 5.9479
1170.3 | 0.0024578 |  | 0.012293 |  | 0.0085595 |  | 0.0098423 |  | 0.011091 |  | 1301.5 | 5.9232
1170.2 | 0.0036851 |  | 0.0073724 |  | 0.011064 |  | 0.014702 |  | 0.0098188 |  | 1301 | 5.9017
1170.1 | 0.0024577 |  | 0.0049161 |  | 0.01599 |  | 0.0036805 |  | 0.0085938 |  | 1300.5 | 5.8777
1170 | 0.0024576 |  | 0.0073768 |  | 0.0061322 |  | 0.0049067 |  | 0.01475 |  | 1300 | 5.8653
1169.9 | 0.013517 |  | 0.0123 |  | 0.011035 |  | -0.0012266 |  | 0.0024567 |  | 1299.5 | 5.8487
1169.8 | 0.0098194 |  | 0.008608 |  | 0.0036882 |  | 0.011046 |  | 0.0012258 |  | 1299 | 5.8205
1169.7 | 0.011111 |  | 0.017229 |  | 0.0049184 |  | -0.0012268 |  | 0.0024536 |  | 1298.5 | 5.8025
1169.6 | 0.012241 |  | 0.0012298 |  | 0.0086029 |  | 0.0098061 |  | 0.0061315 |  | 1298 | 5.779
1169.5 | 0.0086042 |  | 0.011072 |  | 0.0049214 |  | 0.0012315 |  | 0.0073822 |  | 1297.5 | 5.7642
1169.4 | 0.0086132 |  | 0.0073771 |  | 0.013525 |  | 0.0024507 |  | 0.0036811 |  | 1297 | 5.7404
1169.3 | 0.0061523 |  | 0.0049164 |  | 0.0061505 |  | 0.0073594 |  | 0.0049127 |  | 1296.5 | 5.7315
1169.2 | 0.0086093 |  | 0.011071 |  | 0.011059 |  | -0.0024577 |  | 0.012306 |  | 1296 | 5.7118
1169.1 | 0.0098123 |  | 0.0098393 |  | 0.0012254 |  | 0.0086041 |  | 0.012308 |  | 1295.5 | 5.6969
1169 | 0.0024597 |  | 0.017217 |  | 0.012297 |  | 0 |  | 0.0098347 |  | 1295 | 5.6707
1168.9 | 0.0061323 |  | 0.0098514 |  | 0.0073506 |  | 0.0073617 |  | 0.011059 |  | 1294.5 | 5.6597
1168.8 | 0.0098421 |  | 0.0024639 |  | 0.0073823 |  | 0.0073795 |  | 0.011058 |  | 1294 | 5.6458
1168.7 | -0.0012281 |  | 0.011052 |  | 0.011044 |  | 0.0061462 |  | 0.0049182 |  | 1293.5 | 5.6224
1168.6 | 0.0024596 |  | 0.015968 |  | 0.0049161 |  | 0.0049136 |  | 0.0086071 |  | 1293 | 5.6043
1168.5 | 0.0049215 |  | -0.0036808 |  | -0.0012295 |  | 0.0061448 |  | 0.0061498 |  | 1292.5 | 5.5957
1168.4 | 0.0073737 |  | 0.013508 |  | 0.0012282 |  | 0.0049205 |  | 0.002458 |  | 1292 | 5.5832
1168.3 | 0.0012302 |  | 0.012281 |  | 0.011064 |  | -0.0012297 |  | 0.0073817 |  | 1291.5 | 5.5707
1168.2 | 0.0049207 |  | 0.0061385 |  | 0.013526 |  | 0.0012268 |  | 0.0049086 |  | 1291 | 5.5585
1168.1 | -0.0024652 |  | 0.0061368 |  | 0.008621 |  | 0.0049147 |  | 0.0036887 |  | 1290.5 | 5.5441
1168 | 0.0036775 |  | 0.0024617 |  | 0.0012216 |  | -0.0061332 |  | 0.013504 |  | 1290 | 5.5384
1167.9 | 0.0061463 |  | 0.0012287 |  | 0.0049126 |  | 0.0049025 |  | 0.0036806 |  | 1289.5 | 5.5217
1167.8 | 0.011019 |  | 0.003691 |  | 0.0036925 |  | 0.0061478 |  | 0.0024599 |  | 1289 | 5.5197
1167.7 | 0.011075 |  | -0.0024546 |  | 0 |  | 0 |  | 0.01105 |  | 1288.5 | 5.5131
1167.6 | 0.007359 |  | -0.002456 |  | 0.0036867 |  | 0 |  | 0.004916 |  | 1288 | 5.514
1167.5 | 0.0098759 |  | 0.003693 |  | 0.0012274 |  | 0.0049135 |  | 0.0012302 |  | 1287.5 | 5.5117
1167.4 | 0.0073641 |  | 0.0036929 |  | 0.0073661 |  | 0 |  | -0.0086079 |  | 1287 | 5.5
1167.3 | 0.008593 |  | 0.0086267 |  | 0.014724 |  | 0.0024629 |  | 0 |  | 1286.5 | 5.5016
1167.2 | 0.0024605 |  | 0.013499 |  | 0 |  | 0.0024526 |  | 0.0024569 |  | 1286 | 5.5048
1167.1 | 0.0049201 |  | -0.0012287 |  | -0.0024546 |  | 0.0012269 |  | 0.0061432 |  | 1285.5 | 5.513
1167 | 0.0024624 |  | 0.0024615 |  | 0.0049207 |  | 0.0061416 |  | -0.0073794 |  | 1285 | 5.5111
1166.9 | 0 |  | 0.0098331 |  | 0 |  | -0.0037048 |  | -0.0061541 |  | 1284.5 | 5.5248
1166.8 | 0.0024594 |  | -0.0012276 |  | 0.0049104 |  | 0.0073628 |  | 0.0086092 |  | 1284 | 5.5391
1166.7 | 0.0049067 |  | 0.0036991 |  | 0.0024456 |  | 0 |  | 0.0061325 |  | 1283.5 | 5.5478
1166.6 | 0.009835 |  | -0.003684 |  | 0.0036911 |  | 0.0073782 |  | 0.0036817 |  | 1283 | 5.5595
1166.5 | 0.011065 |  | 0.0073746 |  | 0.007399 |  | -0.0024665 |  | -0.002465 |  | 1282.5 | 5.5792
1166.4 | -0.0024593 |  | -0.0012272 |  | -0.0012244 |  | 0.0036852 |  | 0 |  | 1282 | 5.6032
1166.3 | 0 |  | 0.0086009 |  | 0 |  | 0.009812 |  | -0.0024525 |  | 1281.5 | 5.6281
1166.2 | 0.0036855 |  | 0.0086159 |  | 0.0049179 |  | 0.0036845 |  | 0.001227 |  | 1281 | 5.6513
1166.1 | 0.0061421 |  | 0.0049346 |  | 0.001227 |  | 0.0012283 |  | 0.0049148 |  | 1280.5 | 5.6766
1166 | 0.0061508 |  | 0.0024565 |  | 0.0098424 |  | -0.0049232 |  | 0.0012287 |  | 1280 | 5.7108
1165.9 | -0.0036883 |  | 0.012278 |  | -0.0098268 |  | 0.0036828 |  | -0.0036869 |  | 1279.5 | 5.7403
1165.8 | -0.0024587 |  | 0.0036825 |  | 0.0012302 |  | -0.0061506 |  | 0.0024595 |  | 1279 | 5.772
1165.7 | 0.0098399 |  | 0.0086181 |  | -0.0061537 |  | 0.004923 |  | -0.0012304 |  | 1278.5 | 5.8103
1165.6 | 0.0049282 |  | 0.0098305 |  | -0.0024652 |  | 0.0061634 |  | 0.006158 |  | 1278 | 5.8449
1165.5 | 0.0024725 |  | 0.003689 |  | 0.0061542 |  | 0.0086265 |  | 0.0061713 |  | 1277.5 | 5.8892
1165.4 | 0.0061463 |  | 0.0024605 |  | 0.0049072 |  | 2.2183e-07 |  | -0.0012231 |  | 1277 | 5.9412
1165.3 | 0.0036782 |  | 0.0073921 |  | 0.0049228 |  | -0.006097 |  | 0 |  | 1276.5 | 5.9894
1165.2 | 0.0024568 |  | 0.011117 |  | -0.0024494 |  | -0.0024564 |  | -0.007362 |  | 1276 | 6.0476
1165.1 | 0.0049229 |  | -0.0012262 |  | 0.0012321 |  | 0.0012343 |  | -0.0024633 |  | 1275.5 | 6.1044
1165 | 0.014776 |  | 0.010978 |  | -0.0049128 |  | 0.0073861 |  | -0.0024603 |  | 1275 | 6.1523
1164.9 | 0.0098387 |  | 0.0073374 |  | -0.0036843 |  | 0.0036933 |  | -0.0012257 |  | 1274.5 | 6.1982
1164.8 | 0.0085978 |  | -0.0012298 |  | 0.0012276 |  | 0.0061487 |  | 0.0073586 |  | 1274 | 6.2597
1164.7 | 0.0012292 |  | -0.0049276 |  | 0.0036853 |  | 0.0012296 |  | -0.0012297 |  | 1273.5 | 6.2962
1164.6 | 0.0012291 |  | 0.0049201 |  | -0.0085971 |  | 0.0098333 |  | 0.0036844 |  | 1273 | 6.3562
1164.5 | 0.0036862 |  | 0.012287 |  | -0.0098249 |  | -0.004916 |  | 0.0024563 |  | 1272.5 | 6.3781
1164.4 | 0.0036824 |  | -0.001227 |  | -0.0012288 |  | 0.0073644 |  | -0.0049078 |  | 1272 | 6.4548
1164.3 | 0.0073728 |  | -0.0036875 |  | -0.011062 |  | 0.011063 |  | 0.0012284 |  | 1271.5 | 6.4687
1164.2 | 0.0036843 |  | -0.0036848 |  | 0.0024568 |  | 0.0036854 |  | -0.0073781 |  | 1271 | 6.5026
1164.1 | 0.0024574 |  | 0.0049137 |  | -0.0098075 |  | 0.0073741 |  | 0.0012279 |  | 1270.5 | 6.5259
1164 | 0.0012286 |  | 0.0036852 |  | -0.0024564 |  | -0.0086056 |  | 0.006141 |  | 1270 | 6.5535
1163.9 | 0 |  | 0.0012289 |  | -0.0073671 |  | 0.0024591 |  | 0.002458 |  | 1269.5 | 6.6081
1163.8 | 0.0012295 |  | 0.0024587 |  | -0.0049191 |  | 0 |  | -0.0049155 |  | 1269 | 6.5921
1163.7 | 0.0024579 |  | 0.0061441 |  | 0.0012285 |  | 0.0073763 |  | -0.0073743 |  | 1268.5 | 6.6942
1163.6 | 0 |  | 0.0024574 |  | -0.0061546 |  | -0.0024605 |  | 0.0012312 |  | 1268 | 6.6576
1163.5 | 0.0012299 |  | -0.0073786 |  | -0.0073898 |  | -2.23e-07 |  | 0.002461 |  | 1267.5 | 6.7062
1163.4 | 0.001227 |  | -0.0024574 |  | -0.0024523 |  | -0.0012264 |  | 0.0012214 |  | 1267 | 6.7207
1163.3 | 0 |  | 0.0086212 |  | 0 |  | -0.0024562 |  | -0.0097973 |  | 1266.5 | 6.7235
1163.2 | 0.0012275 |  | 0.0024644 |  | 0.0012297 |  | 0.0012277 |  | 0.0036806 |  | 1266 | 6.6723
1163.1 | 2.2336e-07 |  | -0.00859 |  | -0.003688 |  | 0 |  | -0.0012288 |  | 1265.5 | 6.686
1163 | 0.0036841 |  | 0.0012291 |  | 0 |  | 0.0036894 |  | -0.0012287 |  | 1265 | 6.7803
1162.9 | 0.0012277 |  | 0 |  | -0.0073764 |  | 0.002459 |  | 0.0073703 |  | 1264.5 | 6.6846
1162.8 | 0.0012293 |  | -0.0036826 |  | -0.0073705 |  | 0.0049142 |  | -0.0024554 |  | 1264 | 6.7156
1162.7 | 0.0061531 |  | 0.012284 |  | 0 |  | 0.0012303 |  | 0.013518 |  | 1263.5 | 6.7868
1162.6 | 0.0061458 |  | -0.0036864 |  | -0.0012303 |  | 0 |  | 0.0024589 |  | 1263 | 6.7375
1162.5 | 0.0012293 |  | -0.0012283 |  | 0.0049211 |  | -0.0098461 |  | 0 |  | 1262.5 | 6.7311
1162.4 | 0.008604 |  | 0.0024574 |  | -0.0061536 |  | -0.0024619 |  | 0.004918 |  | 1262 | 6.8568
1162.3 | 0.003691 |  | 0.0012283 |  | -0.003687 |  | 0 |  | 0 |  | 1261.5 | 6.8227
1162.2 | 0.0073937 |  | 0.0061441 |  | -0.002459 |  | 0.0036911 |  | 0.0098581 |  | 1261 | 6.9288
1162.1 | 0.0049413 |  | -0.0012278 |  | -0.0061982 |  | 0.00744 |  | -0.0012405 |  | 1260.5 | 6.7467
1162 | 2.2306e-07 |  | 0.0086123 |  | 0.0012238 |  | -0.0012245 |  | 0.006128 |  | 1260 | 6.7925
1161.9 | 0.0012205 |  | 0.0024635 |  | 0.0024387 |  | -0.0048892 |  | 0.0012212 |  | 1259.5 | 6.8464
1161.8 | 0.0061057 |  | 0.0049422 |  | -0.0036701 |  | 0 |  | 0.0024489 |  | 1259 | 6.8032
1161.7 | 0.0012265 |  | 2.2483e-07 |  | 0.0073846 |  | 0.0024553 |  | -0.004913 |  | 1258.5 | 6.8966
1161.6 | 0.0024573 |  | 0.007255 |  | -0.0098228 |  | 0.002458 |  | 0 |  | 1258 | 6.8319
1161.5 | 0.0049178 |  | -0.0036686 |  | -0.0073725 |  | 0.0024614 |  | 0.0012317 |  | 1257.5 | 6.7577
1161.4 | 0.0012296 |  | -0.004919 |  | -2.2373e-07 |  | 0.0024618 |  | -0.0036969 |  | 1257 | 6.8636
1161.3 | -0.0024599 |  | -0.0012306 |  | -0.0061528 |  | 0.007388 |  | -0.0073737 |  | 1256.5 | 6.9139
1161.2 | 0.0049265 |  | 0.0036901 |  | -0.0024672 |  | 0.0037 |  | -0.0012307 |  | 1256 | 6.8902
1161.1 | -0.0024504 |  | 0 |  | 0.0049152 |  | 0.0085816 |  | -0.0036827 |  | 1255.5 | 6.9222
1161 | 0 |  | 0.0036932 |  | 0 |  | 0.0024545 |  | -0.0036846 |  | 1255 | 6.8535
1160.9 | -0.0012286 |  | 0.002465 |  | 0 |  | -0.0049228 |  | -0.0024544 |  | 1254.5 | 6.9615
1160.8 | 0.0024563 |  | 0.0024555 |  | -0.0036864 |  | 0.007368 |  | 0 |  | 1254 | 6.856
1160.7 | 0 |  | 0.0036743 |  | 0.0012278 |  | 0.0024558 |  | 0.0012261 |  | 1253.5 | 6.8411
1160.6 | 0.0036925 |  | 0 |  | 0.0024601 |  | -0.0012301 |  | 0 |  | 1253 | 7.0105
1160.5 | 0.0024636 |  | 0.0085993 |  | 0.0024651 |  | -0.0049322 |  | 0.0012322 |  | 1252.5 | 6.8445
1160.4 | -0.0012249 |  | 0.0024582 |  | 0.0012272 |  | 0.0085983 |  | -0.0073746 |  | 1252 | 6.9048
1160.3 | -0.0024465 |  | -0.0036888 |  | 0.0012291 |  | -0.0073645 |  | 0.0024551 |  | 1251.5 | 6.9633
1160.2 | -0.0049415 |  | 0.0049242 |  | 0.0012388 |  | -0.0037157 |  | 0.0036849 |  | 1251 | 6.9665
1160.1 | 0.001238 |  | -0.0012309 |  | 0 |  | 0.0061514 |  | 0.0074045 |  | 1250.5 | 6.8761
1160 | 0.002423 |  | -0.0024593 |  | -0.0024485 |  | 0.0036689 |  | -0.0024313 |  | 1250 | 6.9706
1159.9 | -0.0036718 |  | 0.0037065 |  | 0 |  | 0.0012269 |  | 0 |  | 1249.5 | 7.0118
1159.8 | 0.0098235 |  | -0.0037055 |  | 0.0024572 |  | -0.0098165 |  | 0.0024593 |  | 1249 | 6.8594
1159.7 | 0.0073783 |  | 0.0048713 |  | 0.0036825 |  | 0.00491 |  | 0.006145 |  | 1248.5 | 7.0192
1159.6 | 0.0012291 |  | 0.0036766 |  | 0.0049249 |  | 0.0036938 |  | 0.0012281 |  | 1248 | 6.98
1159.5 | 0 |  | -0.0036843 |  | 0.002475 |  | 0.0012372 |  | 0.0024684 |  | 1247.5 | 6.9165
1159.4 | 0.0037019 |  | 0.0012289 |  | -0.0036886 |  | 0.0036841 |  | 0 |  | 1247 | 7.0028
1159.3 | 0.0085668 |  | 0.0049195 |  | -0.0012204 |  | 0.0036636 |  | -0.0012133 |  | 1246.5 | 6.982
1159.2 | 0.0036719 |  | 0.001232 |  | 0 |  | 0.0036806 |  | 0.0061383 |  | 1246 | 6.9959
1159.1 | 0.007359 |  | -0.0061814 |  | -0.0012274 |  | 0.0086182 |  | 0.0086055 |  | 1245.5 | 6.9522
1159 | 0.0024583 |  | 0.0073353 |  | 0 |  | 0 |  | 0 |  | 1245 | 7.0049
1158.9 | 0.0012309 |  | -0.0024494 |  | -0.004924 |  | 0 |  | -2.2385e-07 |  | 1244.5 | 6.8463
1158.8 | 0.0012338 |  | 0.006143 |  | 0.0037294 |  | 0.0024613 |  | 0.0074488 |  | 1244 | 6.9718
1158.7 | -0.0024825 |  | -2.2357e-07 |  | 0 |  | 0.0012473 |  | -0.0036258 |  | 1243.5 | 6.944
1158.6 | 0.0059401 |  | 0.0024653 |  | -0.0083367 |  | 0.0023717 |  | -0.0024732 |  | 1243 | 6.9785
1158.5 | -0.0036945 |  | 0.0037122 |  | -0.0012292 |  | 0.0012446 |  | 0.0012198 |  | 1242.5 | 7.0128
1158.4 | 0.0036802 |  | 0.0097565 |  | 0.003684 |  | 0 |  | 0.0073707 |  | 1242 | 7.0958
1158.3 | -0.0024635 |  | 0.0061164 |  | 0.0036912 |  | 0.0012275 |  | 0.0061508 |  | 1241.5 | 6.9591
1158.2 | 0.014777 |  | 0.0085993 |  | 0.0024641 |  | 0.0036902 |  | -0.0061581 |  | 1241 | 7.056
1158.1 | 0 |  | 0.0012271 |  | -0.0024605 |  | 0.0012278 |  | -0.004912 |  | 1240.5 | 6.9535
1158 | -0.0049009 |  | 0.0012285 |  | 0.0012291 |  | -0.008598 |  | 0.0012277 |  | 1240 | 6.9509
1157.9 | 0 |  | 0.006157 |  | -0.0086046 |  | -0.0061483 |  | -0.0024565 |  | 1239.5 | 6.9269
1157.8 | -0.0036915 |  | -0.0049232 |  | 0.0012301 |  | -0.0036913 |  | 0.0036871 |  | 1239 | 6.8943
1157.7 | -0.0012331 |  | 0 |  | 0.0049571 |  | 0 |  | 0 |  | 1238.5 | 6.9918
1157.6 | 0.0012416 |  | 0.0012269 |  | 0.006107 |  | 0.0073104 |  | -0.00123 |  | 1238 | 6.8913
1157.5 | 0 |  | 0.0012292 |  | 0.012228 |  | 0.0012267 |  | -0.0011986 |  | 1237.5 | 6.8801
1157.4 | -0.0024625 |  | 0 |  | 0.01603 |  | 0.0048993 |  | 0.0012366 |  | 1237 | 6.8511
1157.3 | 0.0024328 |  | 0 |  | 0.0049093 |  | -0.0061372 |  | 0.0036779 |  | 1236.5 | 6.8104
1157.2 | 0.011035 |  | 0.0048986 |  | 0 |  | 0 |  | 0.003687 |  | 1236 | 6.8913
1157.1 | 0.0024565 |  | 0.0036441 |  | 0 |  | 0.0036718 |  | -0.0012293 |  | 1235.5 | 6.7994
1157 | 0.0061429 |  | 0.0036767 |  | 0.0036813 |  | 0.0073452 |  | 0.0024517 |  | 1235 | 6.8984
1156.9 | 0.0086246 |  | -0.0024568 |  | 0 |  | 0.0097917 |  | 0.006128 |  | 1234.5 | 6.8366
1156.8 | 0.0048999 |  | 0.0036864 |  | -0.0085736 |  | 0.0024543 |  | 0 |  | 1234 | 6.8934
1156.7 | 0 |  | 0.0036909 |  | 0.0073591 |  | -0.0024539 |  | -0.001228 |  | 1233.5 | 6.9324
1156.6 | -0.0049271 |  | 0 |  | 0.0049201 |  | 0.0024557 |  | 0.002458 |  | 1233 | 6.8455
1156.5 | -0.0024637 |  | -0.0036629 |  | 0 |  | -0.0012276 |  | 0.001229 |  | 1232.5 | 6.711
1156.4 | 0.0012294 |  | -0.0061363 |  | 0 |  | 0.011064 |  | -0.0036863 |  | 1232 | 6.7718
1156.3 | 0.0086471 |  | -0.0036853 |  | 0.0049323 |  | 0.0086644 |  | 0 |  | 1231.5 | 6.7521
1156.2 | 0.0012225 |  | -0.0073803 |  | 0 |  | 0.0024253 |  | 0.0012185 |  | 1231 | 6.7633
1156.1 | 0 |  | -0.0036893 |  | 0.0073575 |  | 0.011051 |  | 0.0024293 |  | 1230.5 | 6.7736
1156 | -0.0036715 |  | -0.0049203 |  | -0.0036783 |  | 0.0012244 |  | -0.0024501 |  | 1230 | 6.6953
1155.9 | -0.0061413 |  | 0.0024626 |  | 0.00246 |  | 0.0049096 |  | 0.0073603 |  | 1229.5 | 6.7095
1155.8 | 0.0012301 |  | 0.0097218 |  | -0.0036918 |  | 0.0036866 |  | 0.0048955 |  | 1229 | 6.6557
1155.7 | 0.0061479 |  | -0.0024539 |  | -0.0049219 |  | -0.0036862 |  | 0.0036903 |  | 1228.5 | 6.6598
1155.6 | 0.0049371 |  | 0 |  | -0.0061084 |  | -0.0049073 |  | 0.01221 |  | 1228 | 6.7039
1155.5 | 0.0050575 |  | -0.0024483 |  | 0.0037566 |  | 0.0050252 |  | -0.0062769 |  | 1227.5 | 6.6271
1155.4 | -0.0024102 |  | -0.0049282 |  | -0.0023988 |  | 0.010766 |  | 0.002419 |  | 1227 | 6.6563
1155.3 | 0.0023614 |  | -0.0024753 |  | -0.0012389 |  | 0.0024694 |  | 0.0061667 |  | 1226.5 | 6.5645
1155.2 | 0.0050073 |  | -0.0037107 |  | 0 |  | 0.0085815 |  | 0.0049143 |  | 1226 | 6.5442
1155.1 | 0.003712 |  | 0.0061371 |  | 0.0024532 |  | -2.2327e-07 |  | 0.006148 |  | 1225.5 | 6.5337
1155 | 0.009717 |  | 0.002383 |  | 0.0061474 |  | -0.004885 |  | 0.0024558 |  | 1225 | 6.4821
1154.9 | 0.0073704 |  | -0.01119 |  | -0.0036876 |  | -0.0012218 |  | -0.0024556 |  | 1224.5 | 6.4522
1154.8 | 0.0048989 |  | 0.0024564 |  | -0.0037043 |  | 0.0012289 |  | -0.0036765 |  | 1224 | 6.4285
1154.7 | -0.0061778 |  | -0.0036746 |  | 0 |  | 0 |  | 0.0049279 |  | 1223.5 | 6.3778
1154.6 | 0.0073249 |  | -0.008601 |  | 0.0036713 |  | 0.0048129 |  | -0.001221 |  | 1223 | 6.3978
1154.5 | 0.0012214 |  | -0.0036757 |  | -0.0086007 |  | -0.0012335 |  | 0.0036471 |  | 1222.5 | 6.2986
1154.4 | 0.0036872 |  | 0.0061153 |  | 0.0073694 |  | 0.012204 |  | 0.0012283 |  | 1222 | 6.305
1154.3 | -0.0061533 |  | 0.0036604 |  | 0.0024722 |  | -0.0049442 |  | -0.0024699 |  | 1221.5 | 6.2722
1154.2 | 0.0048742 |  | -0.0073599 |  | 0.0037072 |  | 0.0085706 |  | 0.003662 |  | 1221 | 6.2322
1154.1 | 0.0049253 |  | -0.0036862 |  | 0.0036861 |  | -0.0048022 |  | 0.0047434 |  | 1220.5 | 6.1814
1154 | -0.0061233 |  | 0 |  | -0.0012126 |  | -0.0024807 |  | 0 |  | 1220 | 6.1482
1153.9 | 0.0024456 |  | -0.003723 |  | -0.0049116 |  | 0.0012056 |  | -0.0012406 |  | 1219.5 | 6.1055
1153.8 | 0.0049286 |  | -0.0023996 |  | 0.0085635 |  | 0.0074895 |  | 0.013495 |  | 1219 | 6.0698
1153.7 | -0.0062469 |  | -0.0061753 |  | 0.0012244 |  | 0.0036941 |  | 0.0023805 |  | 1218.5 | 6.0213
1153.6 | 0.0059406 |  | 0.0049134 |  | 0.0073574 |  | 0.0023895 |  | 0 |  | 1218 | 5.9849
1153.5 | 0.0012455 |  | 0.0062202 |  | -0.0024489 |  | 0.0024951 |  | 0.0097335 |  | 1217.5 | 5.9396
1153.4 | -0.0049026 |  | 0.0012167 |  | 0.0036645 |  | 0.0061175 |  | 0.0036798 |  | 1217 | 5.8895
1153.3 | 0.0073442 |  | 0.001196 |  | -0.0012204 |  | -0.0072813 |  | -0.0036663 |  | 1216.5 | 5.855
1153.2 | 0.0098452 |  | 0.0037421 |  | 0.0036495 |  | -0.0049096 |  | -0.0036824 |  | 1216 | 5.7909
1153.1 | -0.0049204 |  | -0.003662 |  | 0.0012302 |  | 0 |  | 0.0012279 |  | 1215.5 | 5.7509
1153 | 0 |  | 2.2254e-07 |  | 0.0061503 |  | -0.0012283 |  | -0.0036846 |  | 1215 | 5.6945
1152.9 | 0.0024565 |  | -0.0073108 |  | 0.0036849 |  | -0.0036794 |  | -0.014774 |  | 1214.5 | 5.6345
1152.8 | 0 |  | -0.0049257 |  | -0.0036872 |  | -0.013498 |  | 0.0012214 |  | 1214 | 5.5791
1152.7 | -0.0024584 |  | 0.0049231 |  | 0 |  | -0.0061391 |  | 0.0024477 |  | 1213.5 | 5.5234
1152.6 | 0.0061307 |  | -0.0061633 |  | 0.0024562 |  | 0.00488 |  | 0.003719 |  | 1213 | 5.4722
1152.5 | 0.0024876 |  | -0.0098147 |  | 0.0012923 |  | 0.0012828 |  | -0.0012772 |  | 1212.5 | 5.4157
1152.4 | 0.0024964 |  | -0.0061271 |  | 0.0096755 |  | -0.0049147 |  | -0.0012218 |  | 1212 | 5.3471
1152.3 | -0.0024672 |  | -0.0024629 |  | 0.0035816 |  | 0 |  | 0.0011726 |  | 1211.5 | 5.2838
1152.2 | -0.0096331 |  | -0.011031 |  | 0 |  | -0.0036182 |  | 0.0049537 |  | 1211 | 5.221
1152.1 | -0.0046629 |  | -0.0013134 |  | 0.0012178 |  | 0.0012419 |  | 0.0036703 |  | 1210.5 | 5.1514
1152 | 0.0050155 |  | -0.0036346 |  | -0.0085658 |  | -0.0024365 |  | 0.0024161 |  | 1210 | 5.0853
1151.9 | -0.0061934 |  | -0.0094171 |  | -0.0024636 |  | 0 |  | 0.0062004 |  | 1209.5 | 5.0121
1151.8 | -0.0036822 |  | -0.01608 |  | 0.0098077 |  | -0.003728 |  | 0.0012324 |  | 1209 | 4.9381
1151.7 | 0 |  | -0.0024591 |  | 0.001243 |  | -0.0024942 |  | 0.0012146 |  | 1208.5 | 4.8601
1151.6 | 0 |  | -0.0012315 |  | -0.0070774 |  | -0.0047221 |  | 0.001228 |  | 1208 | 4.7844
1151.5 | -0.0049421 |  | -0.006217 |  | 0.0099547 |  | -0.0012476 |  | 2.2103e-07 |  | 1207.5 | 4.7008
1151.4 | 0 |  | -0.0036761 |  | 0 |  | -0.0048743 |  | -0.003687 |  | 1207 | 4.6185
1151.3 | -0.0097388 |  | -0.011797 |  | -0.0036932 |  | -0.0036615 |  | -0.0036834 |  | 1206.5 | 4.5346
1151.2 | -0.0086306 |  | -0.0074397 |  | -0.0086013 |  | -0.015906 |  | 0 |  | 1206 | 4.4383
1151.1 | 0.0024832 |  | -0.0024562 |  | -0.0024892 |  | -0.0061989 |  | -0.0049797 |  | 1205.5 | 4.341
1151 | -0.0036827 |  | -0.0012303 |  | -0.012068 |  | -0.0098319 |  | 0.0024358 |  | 1205 | 4.2407
1150.9 | -0.0082171 |  | -0.0098441 |  | -0.0049408 |  | -0.0085252 |  | 0.0059852 |  | 1204.5 | 4.1361
1150.8 | -0.012546 |  | -0.0060989 |  | -0.0073484 |  | -0.0073669 |  | 0.0024926 |  | 1204 | 4.0295
1150.7 | -0.012358 |  | -0.012459 |  | -0.006164 |  | -0.010961 |  | 2.2332e-07 |  | 1203.5 | 3.9193
1150.6 | -0.0085466 |  | -0.0059948 |  | -0.0012286 |  | -0.0074304 |  | -0.0097242 |  | 1203 | 3.8001
1150.5 | -0.011151 |  | -0.0074675 |  | -0.0083315 |  | -0.0086527 |  | 0.0049396 |  | 1202.5 | 3.6736
1150.4 | -0.0097696 |  | -0.0084354 |  | 0 |  | -0.0083004 |  | -0.0024456 |  | 1202 | 3.5467
1150.3 | -0.0085607 |  | -0.0075437 |  | 0.0012244 |  | -0.0050147 |  | 0.0072402 |  | 1201.5 | 3.4156
1150.2 | 0 |  | 0 |  | -0.0049053 |  | -0.0085981 |  | -0.007438 |  | 1201 | 3.2801
1150.1 | -0.0049109 |  | -0.0061505 |  | 0.0012187 |  | -0.0036808 |  | 0.0035804 |  | 1200.5 | 3.1358
1150 | -0.012239 |  | -0.0024586 |  | -0.0073731 |  | 0.0012288 |  | 0.0062312 |  | 1200 | 2.9828
[11 empty] | 1199.5 | 2.8009
[11 empty] | 1199 | 2.596
[11 empty] | 1198.5 | 2.457
[11 empty] | 1198 | 2.3511
[11 empty] | 1197.5 | 2.3109
[11 empty] | 1197 | 2.2255
[11 empty] | 1196.5 | 2.1271
[11 empty] | 1196 | 2.0164
[11 empty] | 1195.5 | 1.9099
[11 empty] | 1195 | 1.7816
[11 empty] | 1194.5 | 1.6402
[11 empty] | 1194 | 1.4934
[11 empty] | 1193.5 | 1.3126
[11 empty] | 1193 | 1.1138
[11 empty] | 1192.5 | 0.90866
[11 empty] | 1192 | 0.69174
[11 empty] | 1191.5 | 0.4812
[11 empty] | 1191 | 0.29263
[11 empty] | 1190.5 | 0.15059
[11 empty] | 1190 | 0.066708
[11 empty] | 1189.5 | 0.032059
[11 empty] | 1189 | 0.021587
[11 empty] | 1188.5 | 0.016495
[11 empty] | 1188 | 0.014171
[11 empty] | 1187.5 | 0.011884
[11 empty] | 1187 | 0.01001
[11 empty] | 1186.5 | 0.0085156
[11 empty] | 1186 | 0.0087778
[11 empty] | 1185.5 | 0.0076408
[11 empty] | 1185 | 0.0074705
[11 empty] | 1184.5 | 0.0070224
[11 empty] | 1184 | 0.0066359
[11 empty] | 1183.5 | 0.0057241
[11 empty] | 1183 | 0.0059504
[11 empty] | 1182.5 | 0.013887
[11 empty] | 1182 | 0.031316
[11 empty] | 1181.5 | 0.06177
[11 empty] | 1181 | 0.12544
[11 empty] | 1180.5 | 0.23393
[11 empty] | 1180 | 0.38042
[11 empty] | 1179.5 | 0.56198
[11 empty] | 1179 | 0.77173
[11 empty] | 1178.5 | 0.9857
[11 empty] | 1178 | 1.1885
[11 empty] | 1177.5 | 1.3999
[11 empty] | 1177 | 1.5877
[11 empty] | 1176.5 | 1.75
[11 empty] | 1176 | 1.8946
[11 empty] | 1175.5 | 2.0242
[11 empty] | 1175 | 2.1284
[11 empty] | 1174.5 | 2.2415
[11 empty] | 1174 | 2.3364
[11 empty] | 1173.5 | 2.405
[11 empty] | 1173 | 2.4791
[11 empty] | 1172.5 | 2.5834
[11 empty] | 1172 | 2.6805
[11 empty] | 1171.5 | 2.8316
[11 empty] | 1171 | 3.002
[11 empty] | 1170.5 | 3.1608
[11 empty] | 1170 | 3.3105
[11 empty] | 1169.5 | 3.4531
[11 empty] | 1169 | 3.5907
[11 empty] | 1168.5 | 3.7213
[11 empty] | 1168 | 3.8555
[11 empty] | 1167.5 | 3.982
[11 empty] | 1167 | 4.1008
[11 empty] | 1166.5 | 4.2148
[11 empty] | 1166 | 4.3236
[11 empty] | 1165.5 | 4.4281
[11 empty] | 1165 | 4.5338
[11 empty] | 1164.5 | 4.6368
[11 empty] | 1164 | 4.7288
[11 empty] | 1163.5 | 4.8226
[11 empty] | 1163 | 4.9084
[11 empty] | 1162.5 | 4.9951
[11 empty] | 1162 | 5.0766
[11 empty] | 1161.5 | 5.1565
[11 empty] | 1161 | 5.242
[11 empty] | 1160.5 | 5.3083
[11 empty] | 1160 | 5.3844
[11 empty] | 1159.5 | 5.4565
[11 empty] | 1159 | 5.5226
[11 empty] | 1158.5 | 5.5915
[11 empty] | 1158 | 5.6546
[11 empty] | 1157.5 | 5.7178
[11 empty] | 1157 | 5.7779
[11 empty] | 1156.5 | 5.8416
[11 empty] | 1156 | 5.8995
[11 empty] | 1155.5 | 5.9524
[11 empty] | 1155 | 6.0061
[11 empty] | 1154.5 | 6.0583
[11 empty] | 1154 | 6.1251
[11 empty] | 1153.5 | 6.1734
[11 empty] | 1153 | 6.1971
[11 empty] | 1152.5 | 6.2493
[11 empty] | 1152 | 6.3158
[11 empty] | 1151.5 | 6.369
[11 empty] | 1151 | 6.3706
[11 empty] | 1150.5 | 6.4491
[11 empty] | 1150 | 6.486
[11 empty] | 1149.5 | 6.4899
[11 empty] | 1149 | 6.546
[11 empty] | 1148.5 | 6.5491
[11 empty] | 1148 | 6.6383
[11 empty] | 1147.5 | 6.6819
[11 empty] | 1147 | 6.7391
[11 empty] | 1146.5 | 6.7326
[11 empty] | 1146 | 6.7209
[11 empty] | 1145.5 | 6.7335
[11 empty] | 1145 | 6.7487
[11 empty] | 1144.5 | 6.8301
[11 empty] | 1144 | 6.7746
[11 empty] | 1143.5 | 6.9069
[11 empty] | 1143 | 6.9887
[11 empty] | 1142.5 | 6.8791
[11 empty] | 1142 | 7.0046
[11 empty] | 1141.5 | 7.0999
[11 empty] | 1141 | 6.9802
[11 empty] | 1140.5 | 6.9623
[11 empty] | 1140 | 7.0978
[11 empty] | 1139.5 | 7.1262
[11 empty] | 1139 | 7.1709
[11 empty] | 1138.5 | 7.1227
[11 empty] | 1138 | 7.0189
[11 empty] | 1137.5 | 7.3105
[11 empty] | 1137 | 7.2209
[11 empty] | 1136.5 | 7.219
[11 empty] | 1136 | 7.1231
[11 empty] | 1135.5 | 7.1129
[11 empty] | 1135 | 7.083
[11 empty] | 1134.5 | 7.1896
[11 empty] | 1134 | 7.1518
[11 empty] | 1133.5 | 7.0637
[11 empty] | 1133 | 7.0973
[11 empty] | 1132.5 | 7.1265
[11 empty] | 1132 | 7.0924
[11 empty] | 1131.5 | 7.151
[11 empty] | 1131 | 7.1651
[11 empty] | 1130.5 | 7.1859
[11 empty] | 1130 | 7.1306
[11 empty] | 1129.5 | 7.2083
[11 empty] | 1129 | 7.1052
[11 empty] | 1128.5 | 7.1306
[11 empty] | 1128 | 7.3788
[11 empty] | 1127.5 | 7.3143
[11 empty] | 1127 | 7.1795
[11 empty] | 1126.5 | 7.7097
[11 empty] | 1126 | 7.5807
[11 empty] | 1125.5 | 7.0578
[11 empty] | 1125 | 7.3303
[11 empty] | 1124.5 | 7.1189
[11 empty] | 1124 | 7.2643
[11 empty] | 1123.5 | 7.1837
[11 empty] | 1123 | 7.3764
[11 empty] | 1122.5 | 7.2439
[11 empty] | 1122 | 7.2562
[11 empty] | 1121.5 | 7.1517
[11 empty] | 1121 | 7.168
[11 empty] | 1120.5 | 7.1977
[11 empty] | 1120 | 7.3767
[11 empty] | 1119.5 | 7.0869
[11 empty] | 1119 | 7.1271
[11 empty] | 1118.5 | 7.3046
[11 empty] | 1118 | 7.319
[11 empty] | 1117.5 | 7.0744
[11 empty] | 1117 | 7.417
[11 empty] | 1116.5 | 10
[11 empty] | 1116 | 7.3251
[11 empty] | 1115.5 | 7.2617
[11 empty] | 1115 | 7.2304
[11 empty] | 1114.5 | 7.1476
[11 empty] | 1114 | 7.2978
[11 empty] | 1113.5 | 7.3367
[11 empty] | 1113 | 7.2262
[11 empty] | 1112.5 | 7.3233
[11 empty] | 1112 | 7.1946
[11 empty] | 1111.5 | 7.2137
[11 empty] | 1111 | 7.2245
[11 empty] | 1110.5 | 7.1862
[11 empty] | 1110 | 7.3828
[11 empty] | 1109.5 | 7.2539
[11 empty] | 1109 | 7.0569
[11 empty] | 1108.5 | 7.35
[11 empty] | 1108 | 7.3979
[11 empty] | 1107.5 | 7.1724
[11 empty] | 1107 | 7.4262
[11 empty] | 1106.5 | 7.1719
[11 empty] | 1106 | 7.3301
[11 empty] | 1105.5 | 7.1633
[11 empty] | 1105 | 7.114
[11 empty] | 1104.5 | 7.1924
[11 empty] | 1104 | 7.3244
[11 empty] | 1103.5 | 7.1862
[11 empty] | 1103 | 7.2213
[11 empty] | 1102.5 | 7.2466
[11 empty] | 1102 | 7.1007
[11 empty] | 1101.5 | 7.0868
[11 empty] | 1101 | 7.2672
[11 empty] | 1100.5 | 7.2326
[11 empty] | 1100 | 7.3813
[11 empty] | 1099.5 | 7.1342
[11 empty] | 1099 | 7.3164
[11 empty] | 1098.5 | 7.1865
[11 empty] | 1098 | 7.2133
[11 empty] | 1097.5 | 7.3721
[11 empty] | 1097 | 7.0747
[11 empty] | 1096.5 | 7.3382
[11 empty] | 1096 | 7.0909
[11 empty] | 1095.5 | 7.1653
[11 empty] | 1095 | 7.1819
[11 empty] | 1094.5 | 7.3591
[11 empty] | 1094 | 7.1432
[11 empty] | 1093.5 | 6.8801
[11 empty] | 1093 | 7.0508
[11 empty] | 1092.5 | 6.7557
[11 empty] | 1092 | 6.6668
[11 empty] | 1091.5 | 6.6207
[11 empty] | 1091 | 6.4731
[11 empty] | 1090.5 | 6.3404
[11 empty] | 1090 | 6.215
[11 empty] | 1089.5 | 6.1309
[11 empty] | 1089 | 6.0289
[11 empty] | 1088.5 | 5.9421
[11 empty] | 1088 | 5.8573
[11 empty] | 1087.5 | 5.7963
[11 empty] | 1087 | 5.7166
[11 empty] | 1086.5 | 5.6598
[11 empty] | 1086 | 5.603
[11 empty] | 1085.5 | 5.5528
[11 empty] | 1085 | 5.5132
[11 empty] | 1084.5 | 5.4689
[11 empty] | 1084 | 5.443
[11 empty] | 1083.5 | 5.4111
[11 empty] | 1083 | 5.3898
[11 empty] | 1082.5 | 5.3713
[11 empty] | 1082 | 5.3561
[11 empty] | 1081.5 | 5.3465
[11 empty] | 1081 | 5.3429
[11 empty] | 1080.5 | 5.3416
[11 empty] | 1080 | 5.3469
[11 empty] | 1079.5 | 5.3515
[11 empty] | 1079 | 5.3619
[11 empty] | 1078.5 | 5.375
[11 empty] | 1078 | 5.3879
[11 empty] | 1077.5 | 5.4026
[11 empty] | 1077 | 5.4252
[11 empty] | 1076.5 | 5.4484
[11 empty] | 1076 | 5.477
[11 empty] | 1075.5 | 5.5048
[11 empty] | 1075 | 5.533
[11 empty] | 1074.5 | 5.5692
[11 empty] | 1074 | 5.6015
[11 empty] | 1073.5 | 5.6296
[11 empty] | 1073 | 5.6656
[11 empty] | 1072.5 | 5.711
[11 empty] | 1072 | 5.7399
[11 empty] | 1071.5 | 5.7727
[11 empty] | 1071 | 5.8095
[11 empty] | 1070.5 | 5.848
[11 empty] | 1070 | 5.8833
[11 empty] | 1069.5 | 5.9274
[11 empty] | 1069 | 5.9598
[11 empty] | 1068.5 | 6.007
[11 empty] | 1068 | 6.0122
[11 empty] | 1067.5 | 6.0732
[11 empty] | 1067 | 6.0859
[11 empty] | 1066.5 | 6.1272
[11 empty] | 1066 | 6.1529
[11 empty] | 1065.5 | 6.2117
[11 empty] | 1065 | 6.2273
[11 empty] | 1064.5 | 6.2538
[11 empty] | 1064 | 6.3103
[11 empty] | 1063.5 | 6.3347
[11 empty] | 1063 | 6.3657
[11 empty] | 1062.5 | 6.4014
[11 empty] | 1062 | 6.4654
[11 empty] | 1061.5 | 6.48
[11 empty] | 1061 | 6.4395
[11 empty] | 1060.5 | 6.4853
[11 empty] | 1060 | 6.549
[11 empty] | 1059.5 | 6.5489
[11 empty] | 1059 | 6.6518
[11 empty] | 1058.5 | 6.5902
[11 empty] | 1058 | 6.64
[11 empty] | 1057.5 | 6.6507
[11 empty] | 1057 | 6.6664
[11 empty] | 1056.5 | 6.7101
[11 empty] | 1056 | 6.7255
[11 empty] | 1055.5 | 6.6827
[11 empty] | 1055 | 6.6897
[11 empty] | 1054.5 | 6.7617
[11 empty] | 1054 | 6.7574
[11 empty] | 1053.5 | 6.7235
[11 empty] | 1053 | 6.8258
[11 empty] | 1052.5 | 6.8013
[11 empty] | 1052 | 6.7536
[11 empty] | 1051.5 | 6.8605
[11 empty] | 1051 | 6.8239
[11 empty] | 1050.5 | 6.8056
[11 empty] | 1050 | 6.861
[11 empty] | 1049.5 | 6.9929
[11 empty] | 1049 | 6.8825
[11 empty] | 1048.5 | 6.8801
[11 empty] | 1048 | 6.8969
[11 empty] | 1047.5 | 6.9647
[11 empty] | 1047 | 6.8478
[11 empty] | 1046.5 | 6.9084
[11 empty] | 1046 | 6.8206
[11 empty] | 1045.5 | 6.8911
[11 empty] | 1045 | 6.8359
[11 empty] | 1044.5 | 6.8424
[11 empty] | 1044 | 6.8365
[11 empty] | 1043.5 | 6.743
[11 empty] | 1043 | 6.7796
[11 empty] | 1042.5 | 6.6957
[11 empty] | 1042 | 6.6504
[11 empty] | 1041.5 | 6.4997
[11 empty] | 1041 | 6.367
[11 empty] | 1040.5 | 6.2549
[11 empty] | 1040 | 6.1207
[11 empty] | 1039.5 | 6.0406
[11 empty] | 1039 | 5.9464
[11 empty] | 1038.5 | 5.8564
[11 empty] | 1038 | 5.7981
[11 empty] | 1037.5 | 5.7226
[11 empty] | 1037 | 5.6673
[11 empty] | 1036.5 | 5.6084
[11 empty] | 1036 | 5.5641
[11 empty] | 1035.5 | 5.5294
[11 empty] | 1035 | 5.4966
[11 empty] | 1034.5 | 5.4662
[11 empty] | 1034 | 5.4337
[11 empty] | 1033.5 | 5.4215
[11 empty] | 1033 | 5.4138
[11 empty] | 1032.5 | 5.4103
[11 empty] | 1032 | 5.4079
[11 empty] | 1031.5 | 5.4069
[11 empty] | 1031 | 5.4175
[11 empty] | 1030.5 | 5.421
[11 empty] | 1030 | 5.4298
[11 empty] | 1029.5 | 5.4368
[11 empty] | 1029 | 5.4426
[11 empty] | 1028.5 | 5.452
[11 empty] | 1028 | 5.4568
[11 empty] | 1027.5 | 5.4636
[11 empty] | 1027 | 5.4744
[11 empty] | 1026.5 | 5.4778
[11 empty] | 1026 | 5.4773
[11 empty] | 1025.5 | 5.4803
[11 empty] | 1025 | 5.4828
[11 empty] | 1024.5 | 5.4812
[11 empty] | 1024 | 5.4807
[11 empty] | 1023.5 | 5.4754
[11 empty] | 1023 | 5.4749
[11 empty] | 1022.5 | 5.4732
[11 empty] | 1022 | 5.4804
[11 empty] | 1021.5 | 5.4766
[11 empty] | 1021 | 5.4843
[11 empty] | 1020.5 | 5.4853
[11 empty] | 1020 | 5.4991
[11 empty] | 1019.5 | 5.5087
[11 empty] | 1019 | 5.5228
[11 empty] | 1018.5 | 5.53
[11 empty] | 1018 | 5.5419
[11 empty] | 1017.5 | 5.5592
[11 empty] | 1017 | 5.5649
[11 empty] | 1016.5 | 5.5852
[11 empty] | 1016 | 5.5982
[11 empty] | 1015.5 | 5.6175
[11 empty] | 1015 | 5.6266
[11 empty] | 1014.5 | 5.6364
[11 empty] | 1014 | 5.6561
[11 empty] | 1013.5 | 5.6668
[11 empty] | 1013 | 5.6833
[11 empty] | 1012.5 | 5.7032
[11 empty] | 1012 | 5.7165
[11 empty] | 1011.5 | 5.7414
[11 empty] | 1011 | 5.7452
[11 empty] | 1010.5 | 5.7658
[11 empty] | 1010 | 5.7954
[11 empty] | 1009.5 | 5.8174
[11 empty] | 1009 | 5.8358
[11 empty] | 1008.5 | 5.8566
[11 empty] | 1008 | 5.8765
[11 empty] | 1007.5 | 5.9039
[11 empty] | 1007 | 5.9241
[11 empty] | 1006.5 | 5.9594
[11 empty] | 1006 | 5.9841
[11 empty] | 1005.5 | 6.0054
[11 empty] | 1005 | 6.0371
[11 empty] | 1004.5 | 6.0443
[11 empty] | 1004 | 6.0894
[11 empty] | 1003.5 | 6.1003
[11 empty] | 1003 | 6.102
[11 empty] | 1002.5 | 6.1476
[11 empty] | 1002 | 6.1643
[11 empty] | 1001.5 | 6.1573
[11 empty] | 1001 | 6.1811
[11 empty] | 1000.5 | 6.1866
[11 empty] | 1000 | 6.2138